A logo with blue and green letters

AI-generated content may be incorrect.

**Laboratorium Standar Nasional Satuan Ukuran LK-070-IDN**

***Laboratory of National Measurement Standards***

**Sertifikat Kalibrasi**

***Calibration Certificate***

Certificate number: 06122025

Order number: 000012

**Deskripsi Objek yang Dikalibrasi/Diukur**

***Description of object being calibrated or measured***

Jenis alat atau objek : {'id': 'awfdwad', 'en': 'kansjdadk'}

*Type of instrument or object*

Merek/pembuat dan tipe : wereee w43

*Brand/manufacturer and type*

Identifikasi alat :

*Instrument identification*

Nomor seri : 8976543097

*Serial number*

Identifikasi lain : {'id': 'bagus', 'en': 'good'}

*Other identification*

**Identitas Pemilik**

***Owner’s identification***

Nama : BSN SNSU

*Designation*

Alamat : Puspitek 18, Tangerang, Banten

*Address* 18239, ID

**Pengesahan**

***Authorization***

Pejabat yang mengesahkan : Direktur SNSU Mekanika, Radiasi, dan Biologi

*Authorizing officer*

Nama : M

*Name* NIP 11

Tanggal pengesahan : 2025-06-06

*Date of issue (dd/mm/yyyy)*

Jumlah halaman (termasuk halaman ini) :

*Total number of pages including this one*

Dokumen ini disahkan secara elektronik sesuai peraturan yang berlaku dengan sertifikat dari Balai Sertifikasi Elektronik (BsrE) dan tidak memerlukan tanda tangan atau cap. Dokumen asli dapat diperoleh dengan memindai kode QR di samping ini.

*This document is digitally signed. No signature or seal is required. The original document can be obtained by scanning the QR code on the left.*

Kalibrasi atau pengukuran yang dilaporkan dalam sertifikat ini tercakup dalam lingkup akreditasi menurut SNI ISO/IEC 17025 oleh Komite Akreditasi Nasional, kecuali dinyatakan dalam badan sertifikat.

*The calibration or measurement reported in the certificate is covered in the accredication scope according to SNI ISO/IEC 17025 by the National Accredication Committee of Indonesia, unless marked otherwise in the body of certificate.*

Laboratorium Standar Nasional Satuan Ukuran, Badan Standarisasi Nasional (SNSU-BSN)

*Laboratory of National Measurement Standards, Nasional Standardization Agency of Indonesia*

Kompleks Puspiptek, Gedung 420, Setu, Tangerang Selatan, Banten 15314, Indonesia

Telephone +62-21-7560534, +62-21-7560571, Mobile +62-857-8085-7833, Web [www.bsn.go.id](http://www.bsn.go.id), Email nmi@bsn.go.id

A logo with blue and green letters

AI-generated content may be incorrect.No. Sertifikat / *Cert. Number*: 06122025

No. Order / *Order Number*: 000012

Nama Alat/*Instrument Name* : {'id': 'awfdwad', 'en': 'kansjdadk'}

Pembuat/*Manufacturer* : wereee

Model/*Model* : w43

No. Seri/*Serial Number* : 8976543097

Tanggal Kalibrasi/*Calibration Date* : 2025-06-02 – 2025-06-04

Tempat Kalibrasi/*Calibration Place* : Laboratorium SNSU-BSN

**Hasil Kalibrasi/*Calibration Result***

**Kondisi Ruangan/*Environmental Condition***

suhu : (55 {'prefix': '', 'prefix\_pdf': '', 'unit': '\\degreecelsius', 'unit\_pdf': '°C', 'eksponen': '1', 'eksponen\_pdf': '\\tothe{1}'} ± 5) {'prefix': '', 'prefix\_pdf': '', 'unit': '\\degreecelsius', 'unit\_pdf': '°C', 'eksponen': '1', 'eksponen\_pdf': '\\tothe{1}'}

{{ tabel }}

**Catatan/*Notes***

{'id': 'Q', 'en': 'Q'}

{'id': 'bagus', 'en': 'good'}

\frac{1}{2} + \frac{3}{4} = \left|23\right|

12+34=23

22

--------------------------

Statement 1:

id  
en

1\frac{2}{4}

124

gambar di statement

---------------------------

Dikalibrasi oleh/*Calibrated by* : J

(Pelaksana/*Technician*)

Diperiksa oleh/*Checked by* : K

(Penyelia/*Supervisor*)

: L

(Kepala Laboratorium SNSU Biologi)

: M

(Direktur SNSU Mekanika, Radiasi, dan Biologi)

**==== Akhir dari Sertifikat/*End of Certificate* ====**