## Worksheet
[8 empty] | Nomor Order | :
Halaman
Nama Alat |  |  | : |  |  |  |  | Nama Standar | :
Pabrik |  |  | : |  |  |  |  | Pabrik | :
Tipe/No.Seri |  |  | : |  |  |  |  | Tipe/No.Seri | :
Tanggal Kalibrasi |  |  | : |  |  |  |  | Suhu
Tempat Kalibrasi |  |  | : |  |  |  |  | Kelembaban
Tegangan DC |  |  | EMF
Titik Ukur |  | Koneksi |  |  |  | Pembacaan UUC
[6 empty] | 1 | 2 | 3 | 4 | 5 | Unit
20 | mV | Nor |  | 20 | mV
|  | Rev |  | 20 | mV
Tegangan DC
Rentang Alat |  | Penunjukan Standar |  |  |  | Pembacaan UUC
[6 empty] | 1 | 2 | 3 | 4 | 5 | Unit
100 | mV |  | 0 | mV
100 | mV |  | 10 | mV
100 | mV |  | -10 | mV
100 | mV |  | 50 | mV
100 | mV |  | -50 | mV
100 | mV |  | 90 | mV
100 | mV |  | -90 | mV
100 | mV |  | 100 | mV
100 | mV |  | -100 | mV
1 | V |  | 0.1 | V
1 | V |  | -0.1 | V
1 | V |  | 0.5 | V
1 | V |  | -0.5 | V
1 | V |  | 0.9 | V
1 | V |  | -0.9 | V
1 | V |  | 1 | V
1 | V |  | -1 | V
10 | V |  | 1 | V
10 | V |  | -1 | V
10 | V |  | 2 | V
10 | V |  | 3 | V
10 | V |  | 4 | V
10 | V |  | 5 | V
10 | V |  | -5 | V
10 | V |  | 6 | V
10 | V |  | 7 | V
10 | V |  | 8 | V
10 | V |  | 9 | V
10 | V |  | -9 | V
10 | V |  | 10 | V
10 | V |  | -10 | V
100 | V |  | 10 | V
100 | V |  | -10 | V
100 | V |  | 50 | V
100 | V |  | -50 | V
100 | V |  | 90 | V
100 | V |  | -90 | V
100 | V |  | 100 | V
100 | V |  | -100 | V
1000 | V |  | 100 | V
1000 | V |  | -100 | V
1000 | V |  | 500 | V
1000 | V |  | -500 | V
1000 | V |  | 900 | V
1000 | V |  | -900 | V
1000 | V |  | 1000 | V
1000 | V |  | -1000 | V
Arus DC
Rentang Alat |  | Penunjukan Standar |  |  |  | Pembacaan UUC
[6 empty] | 1 | 2 | 3 | 4 | 5 | Unit
100 | µA |  | 0 | µA
100 | µA |  | 10 | µA
100 | µA |  | -10 | µA
100 | µA |  | 50 | µA
100 | µA |  | -50 | µA
100 | µA |  | 90 | µA
100 | µA |  | -90 | µA
100 | µA |  | 100 | µA
100 | µA |  | -100 | µA
1 | mA |  | 0.1 | mA
1 | mA |  | -0.1 | mA
1 | mA |  | 0.5 | mA
1 | mA |  | -0.5 | mA
1 | mA |  | 0.9 | mA
1 | mA |  | -0.9 | mA
1 | mA |  | 1 | mA
1 | mA |  | -1 | mA
10 | mA |  | 1 | mA
10 | mA |  | -1 | mA
10 | mA |  | 2 | mA
10 | mA |  | 3 | mA
10 | mA |  | 4 | mA
10 | mA |  | 5 | mA
10 | mA |  | -5 | mA
10 | mA |  | 6 | mA
10 | mA |  | 7 | mA
10 | mA |  | 8 | mA
10 | mA |  | 9 | mA
10 | mA |  | -9 | mA
10 | mA |  | 10 | mA
10 | mA |  | -10 | mA
100 | mA |  | 10 | mA
100 | mA |  | -10 | mA
100 | mA |  | 50 | mA
100 | mA |  | -50 | mA
100 | mA |  | 90 | mA
100 | mA |  | -90 | mA
100 | mA |  | 100 | mA
100 | mA |  | -100 | mA
1 | A |  | 0.1 | A
1 | A |  | -0.1 | A
1 | A |  | 0.5 | A
1 | A |  | -0.5 | A
1 | A |  | 0.9 | A
1 | A |  | -0.9 | A
1 | A |  | 1 | A
1 | A |  | -1 | A
30 | A |  | 3 | A
30 | A |  | -3 | A
30 | A |  | 15 | A
30 | A |  | -15 | A
30 | A |  | 27 | A
30 | A |  | -27 | A
30 | A |  | 30 | A
30 | A |  | -30 | A
Tegangan AC
Rentang Alat |  | Penunjukan Standar |  |  |  | Pembacaan UUC
[6 empty] | 1 | 2 | 3 | 4 | 5 | Unit
10 | mV | 10 | mV | 20 | Hz
10 | mV | 10 | mV | 50 | Hz
10 | mV | 10 | mV | 1 | kHz
10 | mV | 10 | mV | 10 | kHz
10 | mV | 10 | mV | 20 | kHz
10 | mV | 20 | mV | 50 | kHz
10 | mV | 30 | mV | 100 | kHz
50 | mV | 50 | mV | 40 | Hz
100 | mV | 100 | mV | 80 | Hz
100 | mV | 100 | mV | 20 | Hz
100 | mV | 100 | mV | 50 | Hz
100 | mV | 100 | mV | 1 | kHz
100 | mV | 100 | mV | 10 | kHz
100 | mV | 100 | mV | 20 | kHz
100 | mV | 100 | mV | 50 | kHz
100 | mV | 100 | mV | 100 | kHz
1 | V | 0.1 | V | 20 | Hz
1 | V | 0.1 | V | 50 | Hz
1 | V | 0.1 | V | 1 | kHz
1 | V | 0.1 | V | 10 | kHz
1 | V | 0.1 | V | 20 | kHz
1 | V | 0.1 | V | 50 | kHz
1 | V | 0.1 | V | 100 | kHz
1 | V | 0.3 | V | 80 | Hz
1 | V | 0.4 | V | 100 | Hz
1 | V | 0.5 | V | 160 | Hz
1 | V | 0.5 | V | 200 | Hz
1 | V | 0.6 | V | 400 | Hz
1 | V | 0.6 | V | 500 | Hz
1 | V | 0.8 | V | 1000 | Hz
1 | V | 0.8 | V | 1600 | Hz
1 | V | 1 | V | 20 | Hz
1 | V | 1 | V | 50 | Hz
1 | V | 1 | V | 1 | kHz
1 | V | 1 | V | 10 | kHz
1 | V | 1 | V | 20 | kHz
1 | V | 1 | V | 50 | kHz
1 | V | 1 | V | 100 | kHz
10 | V | 1 | V | 20 | Hz
10 | V | 1 | V | 50 | Hz
10 | V | 1 | V | 1 | kHz
10 | V | 1 | V | 2 | kHz
10 | V | 1 | V | 4 | kHz
10 | V | 1 | V | 5 | kHz
10 | V | 1 | V | 10 | kHz
10 | V | 1 | V | 20 | kHz
10 | V | 1 | V | 50 | kHz
10 | V | 1 | V | 100 | kHz
10 | V | 5 | V | 20 | Hz
10 | V | 5 | V | 50 | Hz
10 | V | 5 | V | 1 | kHz
10 | V | 5 | V | 10 | kHz
10 | V | 5 | V | 20 | kHz
10 | V | 5 | V | 50 | kHz
10 | V | 5 | V | 20 | kHz
10 | V | 10 | V | 20 | Hz
10 | V | 10 | V | 50 | Hz
10 | V | 10 | V | 1 | kHz
10 | V | 10 | V | 10 | kHz
10 | V | 10 | V | 20 | kHz
10 | V | 10 | V | 50 | kHz
10 | V | 10 | V | 100 | kHz
100 | V | 10 | V | 20 | Hz
100 | V | 10 | V | 50 | Hz
100 | V | 10 | V | 1 | kHz
100 | V | 10 | V | 10 | kHz
100 | V | 10 | V | 20 | kHz
100 | V | 10 | V | 50 | kHz
100 | V | 10 | V | 100 | kHz
100 | V | 100 | V | 20 | Hz
100 | V | 100 | V | 50 | Hz
100 | V | 100 | V | 1 | kHz
100 | V | 100 | V | 10 | kHz
100 | V | 100 | V | 20 | kHz
100 | V | 100 | V | 50 | kHz
100 | V | 100 | V | 100 | kHz
1000 | V | 100 | V | 20 | Hz
1000 | V | 100 | V | 50 | Hz
1000 | V | 100 | V | 1 | kHz
1000 | V | 500 | V | 20 | Hz
1000 | V | 500 | V | 50 | Hz
1000 | V | 500 | V | 1 | kHz
1000 | V | 1000 | V | 20 | Hz
1000 | V | 1000 | V | 50 | Hz
1000 | V | 1000 | V | 1 | kHz
Arus AC
Rentang Alat |  | Penunjukan Standar |  |  |  | Pembacaan UUC
[6 empty] | 1 | 2 | 3 | 4 | 5 | Unit
100 | µA | 10 | µA | 20 | Hz
100 | µA | 10 | µA | 50 | Hz
100 | µA | 10 | µA | 1 | kHz
100 | µA | 100 | µA | 20 | Hz
100 | µA | 100 | µA | 50 | Hz
100 | µA | 100 | µA | 1 | kHz
1 | mA | 0.1 | mA | 20 | Hz
1 | mA | 0.1 | mA | 50 | Hz
1 | mA | 0.1 | mA | 1 | kHz
1 | mA | 1 | mA | 20 | Hz
1 | mA | 1 | mA | 50 | Hz
1 | mA | 1 | mA | 1 | kHz
1 | mA | 1 | mA | 5 | kHz
1 | mA | 1 | mA | 10 | kHz
10 | mA | 1 | mA | 20 | Hz
10 | mA | 1 | mA | 50 | Hz
10 | mA | 1 | mA | 1 | kHz
10 | mA | 1 | mA | 5 | kHz
10 | mA | 1 | mA | 10 | kHz
10 | mA | 5 | mA | 20 | Hz
10 | mA | 5 | mA | 50 | Hz
10 | mA | 5 | mA | 1 | kHz
10 | mA | 5 | mA | 5 | kHz
10 | mA | 5 | mA | 10 | kHz
10 | mA | 10 | mA | 20 | Hz
10 | mA | 10 | mA | 50 | Hz
10 | mA | 10 | mA | 1 | kHz
10 | mA | 10 | mA | 5 | kHz
10 | mA | 10 | mA | 10 | kHz
100 | mA | 10 | mA | 20 | Hz
100 | mA | 10 | mA | 50 | Hz
100 | mA | 10 | mA | 1 | kHz
100 | mA | 10 | mA | 5 | kHz
100 | mA | 10 | mA | 10 | kHz
100 | mA | 100 | mA | 20 | Hz
100 | mA | 100 | mA | 50 | Hz
100 | mA | 100 | mA | 1 | kHz
100 | mA | 100 | mA | 5 | kHz
100 | mA | 100 | mA | 10 | kHz
1 | A | 0.1 | A | 20 | Hz
1 | A | 0.1 | A | 50 | Hz
1 | A | 0.1 | A | 1 | kHz
1 | A | 0.1 | A | 5 | kHz
1 | A | 0.1 | A | 10 | kHz
1 | A | 1 | A | 20 | Hz
1 | A | 1 | A | 50 | Hz
1 | A | 1 | A | 1 | kHz
1 | A | 1 | A | 5 | kHz
1 | A | 1 | A | 10 | kHz
30 | A | 3 | A | 20 | Hz
30 | A | 3 | A | 50 | Hz
30 | A | 3 | A | 1 | kHz
30 | A | 15 | A | 20 | Hz
30 | A | 15 | A | 50 | Hz
30 | A | 15 | A | 1 | kHz
30 | A | 30 | A | 20 | Hz
30 | A | 30 | A | 50 | Hz
30 | A | 30 | A | 1 | kHz
Resistansi 2W
Rentang Alat |  | Penunjukan Standar |  |  |  | Pembacaan UUC
[6 empty] | 1 | 2 | 3 | 4 | 5 | Unit
1 | Ω |  | 0 | Ω
1 | Ω |  | 0.1 | Ω
1 | Ω |  | 1 | Ω
10 | Ω |  | 1 | Ω
10 | Ω |  | 10 | Ω
100 | Ω |  | 10 | Ω
100 | Ω |  | 100 | Ω
1 | kΩ |  | 0.1 | kΩ
1 | kΩ |  | 1 | kΩ
10 | kΩ |  | 1 | kΩ
10 | kΩ |  | 10 | kΩ
100 | kΩ |  | 10 | kΩ
100 | kΩ |  | 100 | kΩ
1 | MΩ |  | 0.1 | MΩ
1 | MΩ |  | 1 | MΩ
10 | MΩ |  | 1 | MΩ
10 | MΩ |  | 10 | MΩ
100 | MΩ |  | 10 | MΩ
100 | MΩ |  | 100 | MΩ
1 | GΩ |  | 0.1 | GΩ
1 | GΩ |  | 1 | GΩ

## Data
[8 empty] | Nomor Order | :
Halaman
Nama Alat |  |  | : 8,5 Digit Multimeter |  |  |  |  | Nama Standar | : Multifunction Calibrator
Pabrik |  |  | : Keysight |  |  |  |  | Pabrik | : Fluke
Tipe |  |  | : 3458A |  |  |  |  | Tipe | : 5730A
No.Seri |  |  | : MY12345678 |  |  |  |  | No.Seri | : 4978506
Tanggal Kalibrasi |  |  | : |  |  |  |  | Suhu
Tempat Kalibrasi |  |  | : |  |  |  |  | Kelembaban
Tegangan DC |  |  | EMF
Titik Ukur |  | Koneksi |  |  |  | Pembacaan UUC | [11 empty] | Rata-rata |  | Standar deviasi
[6 empty] | 1 | 2 | 3 | 4 | 5 | Unit
20 | mV | Nor |  | 20 | mV | 20.0006 | 20.0007 | 20.0006 | 20.0007 | 20.0006
|  | Rev |  | 20 | mV | -19.9999 | -20 | -20 | -20.0001 | -19.9999
[6 empty] | -0.00032 | -0.00034 | -0.000315 | -0.000285 | -0.0003 | [7 empty] | 0.000312 | mV | 2.0797e-05
Tegangan DC
Rentang UUC |  | Penunjukan Standar |  |  |  | Pembacaan UUC | [11 empty] | Rata-rata |  | Standar deviasi |  | Resolusi
[6 empty] | 1 | 2 | 3 | 4 | 5 | Unit
100 | mV |  | 0 | mV |  | 0 | 0 | 0 | 0 | 0 | mV | [6 empty] | 0 | mV | 0 | mV | 1e-05 | mV
100 | mV |  | 10 | mV |  | 10 | 10 | 10 | 10 | 10 | mV | [6 empty] | 10 | mV | 0 | mV | 1e-05 | mV
100 | mV |  | -10 | mV |  | -10 | -10 | -10 | -10 | -10 | mV | [6 empty] | -10 | mV | 0 | mV | 1e-05 | mV
100 | mV |  | 50 | mV |  | 50 | 50 | 50 | 50 | 50 | mV | [6 empty] | 50 | mV | 0 | mV | 1e-05 | mV
100 | mV |  | -50 | mV |  | -50 | -50 | -50 | -50 | -50 | mV | [6 empty] | -50 | mV | 0 | mV | 1e-05 | mV
100 | mV |  | 90 | mV |  | 90 | 90 | 90 | 90 | 90 | mV | [6 empty] | 90 | mV | 0 | mV | 1e-05 | mV
100 | mV |  | -90 | mV |  | -90 | -90 | -90 | -90 | -90 | mV | [6 empty] | -90 | mV | 0 | mV | 1e-05 | mV
100 | mV |  | 100 | mV |  | 100 | 100 | 100 | 100 | 100 | mV | [6 empty] | 100 | mV | 0 | mV | 1e-05 | mV
100 | mV |  | -100 | mV |  | -100 | -100 | -100 | -100 | -100 | mV | [6 empty] | -100 | mV | 0 | mV | 1e-05 | mV
1 | V |  | 0.1 | V |  | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | V | [6 empty] | 0.1 | V | 0 | V | 1e-07 | V
1 | V |  | -0.1 | V |  | -0.1 | -0.1 | -0.1 | -0.1 | -0.1 | V | [6 empty] | -0.1 | V | 0 | V | 1e-07 | V
1 | V |  | 0.5 | V |  | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | V | [6 empty] | 0.5 | V | 0 | V | 1e-07 | V
1 | V |  | -0.5 | V |  | -0.5 | -0.5 | -0.5 | -0.5 | -0.5 | V | [6 empty] | -0.5 | V | 0 | V | 1e-07 | V
1 | V |  | 0.9 | V |  | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | V | [6 empty] | 0.9 | V | 0 | V | 1e-07 | V
1 | V |  | -0.9 | V |  | -0.9 | -0.9 | -0.9 | -0.9 | -0.9 | V | [6 empty] | -0.9 | V | 0 | V | 1e-07 | V
1 | V |  | 1 | V |  | 1 | 1 | 1 | 1 | 1 | V | [6 empty] | 1 | V | 0 | V | 1e-07 | V
1 | V |  | -1 | V |  | -1 | -1 | -1 | -1 | -1 | V | [6 empty] | -1 | V | 0 | V | 1e-07 | V
10 | V |  | 1 | V |  | 1 | 1 | 1 | 1 | 1 | V | [6 empty] | 1 | V | 0 | V | 1e-06 | V
10 | V |  | -1 | V |  | -1 | -1 | -1 | -1 | -1 | V | [6 empty] | -1 | V | 0 | V | 1e-06 | V
10 | V |  | 2 | V |  | 2 | 2 | 2 | 2 | 2 | V | [6 empty] | 2 | V | 0 | V | 1e-06 | V
10 | V |  | 3 | V |  | 3 | 3 | 3 | 3 | 3 | V | [6 empty] | 3 | V | 0 | V | 1e-06 | V
10 | V |  | 4 | V |  | 4 | 4 | 4 | 4 | 4 | V | [6 empty] | 4 | V | 0 | V | 1e-06 | V
10 | V |  | 5 | V |  | 5 | 5 | 5 | 5 | 5 | V | [6 empty] | 5 | V | 0 | V | 1e-06 | V
10 | V |  | -5 | V |  | -5 | -5 | -5 | -5 | -5 | V | [6 empty] | -5 | V | 0 | V | 1e-06 | V
10 | V |  | 6 | V |  | 6 | 6 | 6 | 6 | 6 | V | [6 empty] | 6 | V | 0 | V | 1e-06 | V
10 | V |  | 7 | V |  | 7 | 7 | 7 | 7 | 7 | V | [6 empty] | 7 | V | 0 | V | 1e-06 | V
10 | V |  | 8 | V |  | 8 | 8 | 8 | 8 | 8 | V | [6 empty] | 8 | V | 0 | V | 1e-06 | V
10 | V |  | 9 | V |  | 9 | 9 | 9 | 9 | 9 | V | [6 empty] | 9 | V | 0 | V | 1e-06 | V
10 | V |  | -9 | V |  | -9 | -9 | -9 | -9 | -9 | V | [6 empty] | -9 | V | 0 | V | 1e-06 | V
10 | V |  | 10 | V |  | 10 | 10 | 10 | 10 | 10 | V | [6 empty] | 10 | V | 0 | V | 1e-06 | V
10 | V |  | -10 | V |  | -10 | -10 | -10 | -10 | -10 | V | [6 empty] | -10 | V | 0 | V | 1e-06 | V
100 | V |  | 10 | V |  | 10 | 10 | 10 | 10 | 10 | V | [6 empty] | 10 | V | 0 | V | 1e-05 | V
100 | V |  | -10 | V |  | -10 | -10 | -10 | -10 | -10 | V | [6 empty] | -10 | V | 0 | V | 1e-05 | V
100 | V |  | 50 | V |  | 50 | 50 | 50 | 50 | 50 | V | [6 empty] | 50 | V | 0 | V | 1e-05 | V
100 | V |  | -50 | V |  | -50 | -50 | -50 | -50 | -50 | V | [6 empty] | -50 | V | 0 | V | 1e-05 | V
100 | V |  | 90 | V |  | 90 | 90 | 90 | 90 | 90 | V | [6 empty] | 90 | V | 0 | V | 1e-05 | V
100 | V |  | -90 | V |  | -90 | -90 | -90 | -90 | -90 | V | [6 empty] | -90 | V | 0 | V | 1e-05 | V
100 | V |  | 100 | V |  | 100 | 100 | 100 | 100 | 100 | V | [6 empty] | 100 | V | 0 | V | 1e-05 | V
100 | V |  | -100 | V |  | -100 | -100 | -100 | -100 | -100 | V | [6 empty] | -100 | V | 0 | V | 1e-05 | V
1000 | V |  | 100 | V |  | 100 | 100 | 100 | 100 | 100 | V | [6 empty] | 100 | V | 0 | V | 0.0001 | V
1000 | V |  | -100 | V |  | -100 | -100 | -100 | -100 | -100 | V | [6 empty] | -100 | V | 0 | V | 0.0001 | V
1000 | V |  | 500 | V |  | 500 | 500 | 500 | 500 | 500 | V | [6 empty] | 500 | V | 0 | V | 0.0001 | V
1000 | V |  | -500 | V |  | -500 | -500 | -500 | -500 | -500 | V | [6 empty] | -500 | V | 0 | V | 0.0001 | V
1000 | V |  | 900 | V |  | 900 | 900 | 900 | 900 | 900 | V | [6 empty] | 900 | V | 0 | V | 0.0001 | V
1000 | V |  | -900 | V |  | -900 | -900 | -900 | -900 | -900 | V | [6 empty] | -900 | V | 0 | V | 0.0001 | V
1000 | V |  | 1000 | V |  | 1000 | 1000 | 1000 | 1000 | 1000 | V | [6 empty] | 1000 | V | 0 | V | 0.0001 | V
1000 | V |  | -1000 | V |  | -1000 | -1000 | -1000 | -1000 | -1000 | V | [6 empty] | -1000 | V | 0 | V | 0.0001 | V
Arus DC
Rentang UUC |  | Penunjukan Standar |  |  |  | Pembacaan UUC | [11 empty] | Rata-rata |  | Standar deviasi |  | Resolusi
[6 empty] | 1 | 2 | 3 | 4 | 5 | Unit
100 | µA |  | 0 | µA |  | 0 | 0.00202 | 0.00177 | 0.00196 | 0.00182 | µA | [6 empty] | 0.001514 | µA | 0.0008524 | µA | 1e-05 | µA
100 | µA |  | 10 | µA |  | 10 | 9.9982 | 9.9979 | 9.9977 | 9.9982 | µA | [6 empty] | 9.9984 | µA | 0.00092864 | µA | 1e-05 | µA
100 | µA |  | -10 | µA |  | -10 | -10.0016 | -10.0016 | -10.0017 | -10.0014 | µA | [6 empty] | -10.0013 | µA | 0.00071103 | µA | 1e-05 | µA
100 | µA |  | 50 | µA |  | 50 | 49.9969 | 49.9969 | 49.9971 | 49.9972 | µA | [6 empty] | 49.9976 | µA | 0.0013369 | µA | 1e-05 | µA
100 | µA |  | -50 | µA |  | -50 | -50.0005 | -50.0007 | -50.0005 | -50.0007 | µA | [6 empty] | -50.0005 | µA | 0.00027574 | µA | 1e-05 | µA
100 | µA |  | 90 | µA |  | 90 | 89.9961 | 89.9962 | 89.9959 | 89.996 | µA | [6 empty] | 89.9968 | µA | 0.0017855 | µA | 1e-05 | µA
100 | µA |  | -90 | µA |  | -90 | -89.9998 | -89.9995 | -89.9998 | -89.9997 | µA | [6 empty] | -89.9998 | µA | 0.00018042 | µA | 1e-05 | µA
100 | µA |  | 100 | µA |  | 100 | 99.9956 | 99.9958 | 99.9957 | 99.9955 | µA | [6 empty] | 99.9965 | µA | 0.0019438 | µA | 1e-05 | µA
100 | µA |  | -100 | µA |  | -100 | -99.9995 | -99.9994 | -99.9996 | -99.9997 | µA | [6 empty] | -99.9996 | µA | 0.00024652 | µA | 1e-05 | µA
1 | mA |  | 0.1 | mA |  | 0.1 | 0.099989 | 0.09999 | 0.09999 | 0.099991 | mA | [6 empty] | 0.099992 | mA | 4.415e-06 | mA | 1e-07 | mA
1 | mA |  | -0.1 | mA |  | -0.1 | -0.10002 | -0.10002 | -0.10002 | -0.10002 | mA | [6 empty] | -0.10002 | mA | 1.0205e-05 | mA | 1e-07 | mA
1 | mA |  | 0.5 | mA |  | 0.5 | 0.50002 | 0.50003 | 0.50002 | 0.50003 | mA | [6 empty] | 0.50002 | mA | 1.1335e-05 | mA | 1e-07 | mA
1 | mA |  | -0.5 | mA |  | -0.5 | -0.50005 | -0.50005 | -0.50005 | -0.50005 | mA | [6 empty] | -0.50004 | mA | 2.3564e-05 | mA | 1e-07 | mA
1 | mA |  | 0.9 | mA |  | 0.9 | 0.90006 | 0.90006 | 0.90005 | 0.90006 | mA | [6 empty] | 0.90004 | mA | 2.5123e-05 | mA | 1e-07 | mA
1 | mA |  | -0.9 | mA |  | -0.9 | -0.90008 | -0.90008 | -0.90008 | -0.90008 | mA | [6 empty] | -0.90007 | mA | 3.6664e-05 | mA | 1e-07 | mA
1 | mA |  | 1 | mA |  | 1 | 1.0001 | 1.0001 | 1.0001 | 1.0001 | mA | [6 empty] | 1.0001 | mA | 2.8424e-05 | mA | 1e-07 | mA
1 | mA |  | -1 | mA |  | -1 | -1.0001 | -1.0001 | -1.0001 | -1.0001 | mA | [6 empty] | -1.0001 | mA | 4.0364e-05 | mA | 1e-07 | mA
10 | mA |  | 1 | mA |  | 1 | 0.99985 | 0.99987 | 0.99983 | 0.99986 | mA | [6 empty] | 0.99988 | mA | 6.7833e-05 | mA | 1e-06 | mA
10 | mA |  | -1 | mA |  | -1 | -1.0002 | -1.0002 | -1.0002 | -1.0002 | mA | [6 empty] | -1.0001 | mA | 8.2415e-05 | mA | 1e-06 | mA
10 | mA |  | 2 | mA |  | 2 | 1.9999 | 1.9999 | 1.9999 | 1.9999 | mA | [6 empty] | 1.9999 | mA | 6.0219e-05 | mA | 1e-06 | mA
10 | mA |  | 3 | mA |  | 3 | 2.9999 | 2.9999 | 2.9999 | 2.9999 | mA | [6 empty] | 2.9999 | mA | 2.9115e-05 | mA | 1e-06 | mA
10 | mA |  | 4 | mA |  | 4 | 4 | 4 | 4 | 4 | mA | [6 empty] | 4 | mA | 1.8069e-05 | mA | 1e-06 | mA
10 | mA |  | 5 | mA |  | 5 | 5 | 5 | 5 | 5 | mA | [6 empty] | 5 | mA | 1.3236e-05 | mA | 1e-06 | mA
10 | mA |  | -5 | mA |  | -5 | -5.0003 | -5.0003 | -5.0003 | -5.0003 | mA | [6 empty] | -5.0002 | mA | 0.00011765 | mA | 1e-06 | mA
10 | mA |  | 6 | mA |  | 6 | 6 | 6 | 6 | 6 | mA | [6 empty] | 6 | mA | 1.1371e-05 | mA | 1e-06 | mA
10 | mA |  | 7 | mA |  | 7 | 7 | 7 | 7 | 7.0001 | mA | [6 empty] | 7 | mA | 2.1371e-05 | mA | 1e-06 | mA
10 | mA |  | 8 | mA |  | 8 | 8 | 8.0001 | 8.0001 | 8.0001 | mA | [6 empty] | 8 | mA | 2.842e-05 | mA | 1e-06 | mA
10 | mA |  | 9 | mA |  | 9 | 9.0001 | 9.0001 | 9.0001 | 9.0001 | mA | [6 empty] | 9.0001 | mA | 3.8318e-05 | mA | 1e-06 | mA
10 | mA |  | -9 | mA |  | -9 | -9.0004 | -9.0004 | -9.0004 | -9.0003 | mA | [6 empty] | -9.0003 | mA | 0.00016201 | mA | 1e-06 | mA
10 | mA |  | 10 | mA |  | 10 | 10.0001 | 10.0001 | 10.0001 | 10.0001 | mA | [6 empty] | 10.0001 | mA | 5.2101e-05 | mA | 1e-06 | mA
10 | mA |  | -10 | mA |  | -10 | -10.0004 | -10.0004 | -10.0004 | -10.0004 | mA | [6 empty] | -10.0003 | mA | 0.00017802 | mA | 1e-06 | mA
100 | mA |  | 10 | mA |  | 10 | 10.0001 | 10.0001 | 10.0002 | 10.0003 | mA | [6 empty] | 10.0002 | mA | 0.0001348 | mA | 1e-05 | mA
100 | mA |  | -10 | mA |  | -10 | -10.0035 | -10.0035 | -10.0033 | -10.0033 | mA | [6 empty] | -10.0027 | mA | 0.0015187 | mA | 1e-05 | mA
100 | mA |  | 50 | mA |  | 50 | 50.0078 | 50.0079 | 50.0079 | 50.0078 | mA | [6 empty] | 50.0063 | mA | 0.0035067 | mA | 1e-05 | mA
100 | mA |  | -50 | mA |  | -50 | -50.0103 | -50.0103 | -50.0104 | -50.0106 | mA | [6 empty] | -50.0083 | mA | 0.0046446 | mA | 1e-05 | mA
100 | mA |  | 90 | mA |  | 90 | 90.0155 | 90.0153 | 90.0154 | 90.0153 | mA | [6 empty] | 90.0123 | mA | 0.0068673 | mA | 1e-05 | mA
100 | mA |  | -90 | mA |  | -90 | -90.018 | -90.0183 | -90.0182 | -90.0182 | mA | [6 empty] | -90.0145 | mA | 0.0081309 | mA | 1e-05 | mA
100 | mA |  | 100 | mA |  | 100 | 100.0174 | 100.0172 | 100.0172 | 100.0173 | mA | [6 empty] | 100.0138 | mA | 0.0077237 | mA | 1e-05 | mA
100 | mA |  | -100 | mA |  | -100 | -100.02 | -100.0203 | -100.0202 | -100.0201 | mA | [6 empty] | -100.0161 | mA | 0.0090243 | mA | 1e-05 | mA
1 | A |  | 0.1 | A |  | 0.1 | 0.099979 | 0.09998 | 0.099982 | 0.099979 | A | [6 empty] | 0.099984 | A | 9.0323e-06 | A | 1e-07 | A
1 | A |  | -0.1 | A |  | -0.1 | -0.10002 | -0.10002 | -0.10002 | -0.10002 | A | [6 empty] | -0.10001 | A | 8.0117e-06 | A | 1e-07 | A
1 | A |  | 0.5 | A |  | 0.5 | 0.5 | 0.49996 | 0.5 | 0.49999 | A | [6 empty] | 0.49999 | A | 1.7233e-05 | A | 1e-07 | A
1 | A |  | -0.5 | A |  | -0.5 | -0.50001 | -0.50003 | -0.50001 | -0.50002 | A | [6 empty] | -0.50001 | A | 1.0999e-05 | A | 1e-07 | A
1 | A |  | 0.9 | A |  | 0.9 | 0.89996 | 0.89995 | 0.89997 | 0.89997 | A | [6 empty] | 0.89997 | A | 1.7921e-05 | A | 1e-07 | A
1 | A |  | -0.9 | A |  | -0.9 | -0.9 | -0.9 | -0.90001 | -0.89999 | A | [6 empty] | -0.9 | A | 4.8174e-06 | A | 1e-07 | A
1 | A |  | 1 | A |  | 1 | 0.99997 | 0.99997 | 0.99996 | 0.99996 | A | Rentang Pembacaan 8508 / 3458 |  | Titik Ukur Tegangan 5730 |  |  |  | 0.99997 | A | 1.6529e-05 | A | 1e-07 | A
1 | A |  | -1 | A |  | -1 | -0.99997 | -0.99999 | -0.99997 | -1 | A | [6 empty] | -0.99998 | A | 1.6621e-05 | A | 1e-07 | A
30 | A |  | 3 | A | STD (8508) | 30 | 29.9477 | 29.9472 | 29.9464 | 29.9467 | mV | 100 | mV | 0.3 | V |  |  | 29.9576 | mV | 0.023703 | mV | 1e-05 | mV
[5 empty] | UUC (3458) | 3 | 2.9937 | 2.9939 | 2.9937 | 2.9939 | mV | 100 | mV |  |  |  |  | 2.995 | mV | 0.0027756 | mV | 1e-05 | mV
30 | A |  | -3 | A | STD (8508) | -30 | -29.9914 | -29.9907 | -29.991 | -29.9916 | mV | 100 | mV | -0.3 | V |  |  | -29.9929 | mV | 0.0039582 | mV | 1e-05 | mV
[5 empty] | UUC (3458) | -3 | -2.9982 | -2.9981 | -2.9981 | -2.9981 | mV | 100 | mV |  |  |  |  | -2.9985 | mV | 0.00084716 | mV | 1e-05 | mV
30 | A |  | 15 | A | STD (8508) | 15 | 14.9784 | 14.9786 | 14.9784 | 14.9784 | mV | 100 | mV | 1.5 | V |  |  | 14.9828 | mV | 0.00964 | mV | 1e-05 | mV
[5 empty] | UUC (3458) | 15 | 14.9777 | 14.9776 | 14.9778 | 14.978 | mV | 100 | mV |  |  |  |  | 14.9822 | mV | 0.0099363 | mV | 1e-05 | mV
30 | A |  | -15 | A | STD (8508) | -15 | -14.9841 | -14.9842 | -14.9842 | -14.9843 | mV | 100 | mV | -1.5 | V |  |  | -14.9874 | mV | 0.0070642 | mV | 1e-05 | mV
[5 empty] | UUC (3458) | -15 | -14.9798 | -14.9796 | -14.9797 | -14.9798 | mV | 100 | mV |  |  |  |  | -14.9838 | mV | 0.0090576 | mV | 1e-05 | mV
30 | A |  | 27 | A | STD (8508) | 27 | 26.9567 | 26.957 | 26.9568 | 26.9568 | mV | 100 | mV | 0.27 | V |  |  | 26.9655 | mV | 0.019312 | mV | 1e-05 | mV
[5 empty] | UUC (3458) | 27 | 26.9535 | 26.9534 | 26.9532 | 26.9534 | mV | 100 | mV |  |  |  |  | 26.9627 | mV | 0.020849 | mV | 1e-05 | mV
30 | A |  | -27 | A | STD (8508) | -27 | -26.9743 | -26.9741 | -26.9742 | -26.974 | mV | 100 | mV | -0.27 | V |  |  | -26.9793 | mV | 0.011554 | mV | 1e-05 | mV
[5 empty] | UUC (3458) | -27 | -26.966 | -26.9661 | -26.9659 | -26.9658 | mV | 100 | mV |  |  |  |  | -26.9728 | mV | 0.015227 | mV | 1e-05 | mV
30 | A |  | 30 | A | STD (8508) | 30 | 29.9533 | 29.9535 | 29.9536 | 29.9534 | mV | 100 | mV | 0.3 | V |  |  | 29.9628 | mV | 0.020819 | mV | 1e-05 | mV
[5 empty] | UUC (3458) | 30 | 29.9472 | 29.9471 | 29.9479 | 29.9467 | mV | 100 | mV |  |  |  |  | 29.9578 | mV | 0.023609 | mV | 1e-05 | mV
30 | A |  | -30 | A | STD (8508) | -30 | -29.9704 | -29.9703 | -29.9703 | -29.9702 | mV | 100 | mV | -0.3 | V |  |  | -29.9762 | mV | 0.013279 | mV | 1e-05 | mV
[5 empty] | UUC (3458) | -30 | -29.9595 | -29.9591 | -29.9589 | -29.9586 | mV | 100 | mV |  |  |  |  | -29.9672 | mV | 0.018329 | mV | 1e-05 | mV
Tegangan AC
Rentang UUC |  | Penunjukan Standar |  |  |  | Pembacaan UUC | [11 empty] | Rata-rata |  | Standar deviasi |  | Resolusi
[6 empty] | 1 | 2 | 3 | 4 | 5 | Unit
10 | mV | 10 | mV | 20 | Hz | 10 | 10 | 10 | 10 | 10 | mV | [6 empty] | 10 | mV | 0 | mV | 1e-05 | mV
10 | mV | 10 | mV | 50 | Hz | 10 | 10 | 10 | 10 | 10 | mV | [6 empty] | 10 | mV | 0 | mV | 1e-05 | mV
10 | mV | 10 | mV | 1 | kHz | 10 | 10 | 10 | 10 | 10 | mV | [6 empty] | 10 | mV | 0 | mV | 1e-05 | mV
10 | mV | 10 | mV | 10 | kHz | 10 | 10 | 10 | 10 | 10 | mV | [6 empty] | 10 | mV | 0 | mV | 1e-05 | mV
10 | mV | 10 | mV | 20 | kHz | 10 | 10 | 10 | 10 | 10 | mV | [6 empty] | 10 | mV | 0 | mV | 1e-05 | mV
100 | mV | 20 | mV | 50 | kHz | 20 | 20 | 20 | 20 | 20 | mV | [6 empty] | 20 | mV | 0 | mV | 0.0001 | mV
100 | mV | 30 | mV | 100 | kHz | 30 | 30 | 30 | 30 | 30 | mV | [6 empty] | 30 | mV | 0 | mV | 0.0001 | mV
100 | mV | 50 | mV | 40 | Hz | 50 | 50 | 50 | 50 | 50 | mV | [6 empty] | 50 | mV | 0 | mV | 0.0001 | mV
100 | mV | 100 | mV | 80 | Hz | 100 | 100 | 100 | 100 | 100 | mV | [6 empty] | 100 | mV | 0 | mV | 0.0001 | mV
100 | mV | 100 | mV | 20 | Hz | 100 | 100 | 100 | 100 | 100 | mV | [6 empty] | 100 | mV | 0 | mV | 0.0001 | mV
100 | mV | 100 | mV | 50 | Hz | 100 | 100 | 100 | 100 | 100 | mV | [6 empty] | 100 | mV | 0 | mV | 0.0001 | mV
100 | mV | 100 | mV | 1 | kHz | 100 | 100 | 100 | 100 | 100 | mV | [6 empty] | 100 | mV | 0 | mV | 0.0001 | mV
100 | mV | 100 | mV | 10 | kHz | 100 | 100 | 100 | 100 | 100 | mV | [6 empty] | 100 | mV | 0 | mV | 0.0001 | mV
100 | mV | 100 | mV | 20 | kHz | 100 | 100 | 100 | 100 | 100 | mV | [6 empty] | 100 | mV | 0 | mV | 0.0001 | mV
100 | mV | 100 | mV | 50 | kHz | 100 | 100 | 100 | 100 | 100 | mV | [6 empty] | 100 | mV | 0 | mV | 0.0001 | mV
100 | mV | 100 | mV | 100 | kHz | 100 | 100 | 100 | 100 | 100 | mV | [6 empty] | 100 | mV | 0 | mV | 0.0001 | mV
1 | V | 0.1 | V | 20 | Hz | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | V | [6 empty] | 0.1 | V | 0 | V | 1e-06 | V
1 | V | 0.1 | V | 50 | Hz | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | V | [6 empty] | 0.1 | V | 0 | V | 1e-06 | V
1 | V | 0.1 | V | 1 | kHz | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | V | [6 empty] | 0.1 | V | 0 | V | 1e-06 | V
1 | V | 0.1 | V | 10 | kHz | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | V | [6 empty] | 0.1 | V | 0 | V | 1e-06 | V
1 | V | 0.1 | V | 20 | kHz | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | V | [6 empty] | 0.1 | V | 0 | V | 1e-06 | V
1 | V | 0.1 | V | 50 | kHz | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | V | [6 empty] | 0.1 | V | 0 | V | 1e-06 | V
1 | V | 0.1 | V | 100 | kHz | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | V | [6 empty] | 0.1 | V | 0 | V | 1e-06 | V
1 | V | 0.3 | V | 80 | Hz | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | V | [6 empty] | 0.3 | V | 0 | V | 1e-06 | V
1 | V | 0.4 | V | 100 | Hz | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | V | [6 empty] | 0.4 | V | 0 | V | 1e-06 | V
1 | V | 0.5 | V | 160 | Hz | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | V | [6 empty] | 0.5 | V | 0 | V | 1e-06 | V
1 | V | 0.5 | V | 200 | Hz | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | V | [6 empty] | 0.5 | V | 0 | V | 1e-06 | V
1 | V | 0.6 | V | 400 | Hz | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | V | [6 empty] | 0.6 | V | 0 | V | 1e-06 | V
1 | V | 0.6 | V | 500 | Hz | 0.6 | 0.6 | 0.6 | 0.6 | 0.6 | V | [6 empty] | 0.6 | V | 0 | V | 1e-06 | V
1 | V | 0.8 | V | 1000 | Hz | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | V | [6 empty] | 0.8 | V | 0 | V | 1e-06 | V
1 | V | 0.8 | V | 1600 | Hz | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | V | [6 empty] | 0.8 | V | 0 | V | 1e-06 | V
1 | V | 1 | V | 20 | Hz | 1 | 1 | 1 | 1 | 1 | V | [6 empty] | 1 | V | 0 | V | 1e-06 | V
1 | V | 1 | V | 50 | Hz | 1 | 1 | 1 | 1 | 1 | V | [6 empty] | 1 | V | 0 | V | 1e-06 | V
1 | V | 1 | V | 1 | kHz | 1 | 1 | 1 | 1 | 1 | V | [6 empty] | 1 | V | 0 | V | 1e-06 | V
1 | V | 1 | V | 10 | kHz | 1 | 1 | 1 | 1 | 1 | V | [6 empty] | 1 | V | 0 | V | 1e-06 | V
1 | V | 1 | V | 20 | kHz | 1 | 1 | 1 | 1 | 1 | V | [6 empty] | 1 | V | 0 | V | 1e-06 | V
1 | V | 1 | V | 50 | kHz | 1 | 1 | 1 | 1 | 1 | V | [6 empty] | 1 | V | 0 | V | 1e-06 | V
1 | V | 1 | V | 100 | kHz | 1 | 1 | 1 | 1 | 1 | V | [6 empty] | 1 | V | 0 | V | 1e-06 | V
10 | V | 1 | V | 20 | Hz | 1 | 1 | 1 | 1 | 1 | V | [6 empty] | 1 | V | 0 | V | 1e-05 | V
10 | V | 1 | V | 50 | Hz | 1 | 1 | 1 | 1 | 1 | V | [6 empty] | 1 | V | 0 | V | 1e-05 | V
10 | V | 1 | V | 1 | kHz | 1 | 1 | 1 | 1 | 1 | V | [6 empty] | 1 | V | 0 | V | 1e-05 | V
10 | V | 1 | V | 2 | kHz | 1 | 1 | 1 | 1 | 1 | V | [6 empty] | 1 | V | 0 | V | 1e-05 | V
10 | V | 1 | V | 4 | kHz | 1 | 1 | 1 | 1 | 1 | V | [6 empty] | 1 | V | 0 | V | 1e-05 | V
10 | V | 1 | V | 5 | kHz | 1 | 1 | 1 | 1 | 1 | V | [6 empty] | 1 | V | 0 | V | 1e-05 | V
10 | V | 1 | V | 10 | kHz | 1 | 1 | 1 | 1 | 1 | V | [6 empty] | 1 | V | 0 | V | 1e-05 | V
10 | V | 1 | V | 20 | kHz | 1 | 1 | 1 | 1 | 1 | V | [6 empty] | 1 | V | 0 | V | 1e-05 | V
10 | V | 1 | V | 50 | kHz | 1 | 1 | 1 | 1 | 1 | V | [6 empty] | 1 | V | 0 | V | 1e-05 | V
10 | V | 1 | V | 100 | kHz | 1 | 1 | 1 | 1 | 1 | V | [6 empty] | 1 | V | 0 | V | 1e-05 | V
10 | V | 5 | V | 20 | Hz | 5 | 5 | 5 | 5 | 5 | V | [6 empty] | 5 | V | 0 | V | 1e-05 | V
10 | V | 5 | V | 50 | Hz | 5 | 5 | 5 | 5 | 5 | V | [6 empty] | 5 | V | 0 | V | 1e-05 | V
10 | V | 5 | V | 1 | kHz | 5 | 5 | 5 | 5 | 5 | V | [6 empty] | 5 | V | 0 | V | 1e-05 | V
10 | V | 5 | V | 10 | kHz | 5 | 5 | 5 | 5 | 5 | V | [6 empty] | 5 | V | 0 | V | 1e-05 | V
10 | V | 5 | V | 20 | kHz | 5 | 5 | 5 | 5 | 5 | V | [6 empty] | 5 | V | 0 | V | 1e-05 | V
10 | V | 5 | V | 50 | kHz | 5 | 5 | 5 | 5 | 5 | V | [6 empty] | 5 | V | 0 | V | 1e-05 | V
10 | V | 5 | V | 20 | kHz | 5 | 5 | 5 | 5 | 5 | V | [6 empty] | 5 | V | 0 | V | 1e-05 | V
10 | V | 10 | V | 20 | Hz | 10 | 10 | 10 | 10 | 10 | V | [6 empty] | 10 | V | 0 | V | 1e-05 | V
10 | V | 10 | V | 50 | Hz | 10 | 10 | 10 | 10 | 10 | V | [6 empty] | 10 | V | 0 | V | 1e-05 | V
10 | V | 10 | V | 1 | kHz | 10 | 10 | 10 | 10 | 10 | V | [6 empty] | 10 | V | 0 | V | 1e-05 | V
10 | V | 10 | V | 10 | kHz | 10 | 10 | 10 | 10 | 10 | V | [6 empty] | 10 | V | 0 | V | 1e-05 | V
10 | V | 10 | V | 20 | kHz | 10 | 10 | 10 | 10 | 10 | V | [6 empty] | 10 | V | 0 | V | 1e-05 | V
10 | V | 10 | V | 50 | kHz | 10 | 10 | 10 | 10 | 10 | V | [6 empty] | 10 | V | 0 | V | 1e-05 | V
10 | V | 10 | V | 100 | kHz | 10 | 10 | 10 | 10 | 10 | V | [6 empty] | 10 | V | 0 | V | 1e-05 | V
100 | V | 10 | V | 20 | Hz | 10 | 10 | 10 | 10 | 10 | V | [6 empty] | 10 | V | 0 | V | 0.0001 | V
100 | V | 10 | V | 50 | Hz | 10 | 10 | 10 | 10 | 10 | V | [6 empty] | 10 | V | 0 | V | 0.0001 | V
100 | V | 10 | V | 1 | kHz | 10 | 10 | 10 | 10 | 10 | V | [6 empty] | 10 | V | 0 | V | 0.0001 | V
100 | V | 10 | V | 10 | kHz | 10 | 10 | 10 | 10 | 10 | V | [6 empty] | 10 | V | 0 | V | 0.0001 | V
100 | V | 10 | V | 20 | kHz | 10 | 10 | 10 | 10 | 10 | V | [6 empty] | 10 | V | 0 | V | 0.0001 | V
100 | V | 10 | V | 50 | kHz | 10 | 10 | 10 | 10 | 10 | V | [6 empty] | 10 | V | 0 | V | 0.0001 | V
100 | V | 10 | V | 100 | kHz | 10 | 10 | 10 | 10 | 10 | V | [6 empty] | 10 | V | 0 | V | 0.0001 | V
100 | V | 100 | V | 20 | Hz | 100 | 100 | 100 | 100 | 100 | V | [6 empty] | 100 | V | 0 | V | 0.0001 | V
100 | V | 100 | V | 50 | Hz | 100 | 100 | 100 | 100 | 100 | V | [6 empty] | 100 | V | 0 | V | 0.0001 | V
100 | V | 100 | V | 1 | kHz | 100 | 100 | 100 | 100 | 100 | V | [6 empty] | 100 | V | 0 | V | 0.0001 | V
100 | V | 100 | V | 10 | kHz | 100 | 100 | 100 | 100 | 100 | V | [6 empty] | 100 | V | 0 | V | 0.0001 | V
100 | V | 100 | V | 20 | kHz | 100 | 100 | 100 | 100 | 100 | V | [6 empty] | 100 | V | 0 | V | 0.0001 | V
100 | V | 100 | V | 50 | kHz | 100 | 100 | 100 | 100 | 100 | V | [6 empty] | 100 | V | 0 | V | 0.0001 | V
100 | V | 100 | V | 100 | kHz | 100 | 100 | 100 | 100 | 100 | V | [6 empty] | 100 | V | 0 | V | 0.0001 | V
1000 | V | 100 | V | 20 | Hz | 100 | 100 | 100 | 100 | 100 | V | [6 empty] | 100 | V | 0 | V | 0.001 | V
1000 | V | 100 | V | 50 | Hz | 100 | 100 | 100 | 100 | 100 | V | [6 empty] | 100 | V | 0 | V | 0.001 | V
1000 | V | 100 | V | 1 | kHz | 100 | 100 | 100 | 100 | 100 | V | [6 empty] | 100 | V | 0 | V | 0.001 | V
1000 | V | 500 | V | 20 | Hz | 500 | 500 | 500 | 500 | 500 | V | [6 empty] | 500 | V | 0 | V |  | V
1000 | V | 500 | V | 50 | Hz | 500 | 500 | 500 | 500 | 500 | V | [6 empty] | 500 | V | 0 | V | 0.001 | V
1000 | V | 500 | V | 1 | kHz | 500 | 500 | 500 | 500 | 500 | V | [6 empty] | 500 | V | 0 | V | 0.001 | V
1000 | V | 900 | V | 20 | Hz | 900 | 900 | 900 | 900 | 900 | V | [6 empty] | 900 | V | 0 | V |  | V
1000 | V | 900 | V | 50 | Hz | 900 | 900 | 900 | 900 | 900 | V | [6 empty] | 900 | V | 0 | V | 0.001 | V
1000 | V | 900 | V | 1 | kHz | 900 | 900 | 900 | 900 | 900 | V | [6 empty] | 900 | V | 0 | V | 0.001 | V
Arus AC
Rentang UUC |  | Penunjukan Standar |  |  |  | Pembacaan UUC | [11 empty] | Rata-rata |  | Standar deviasi |  | Resolusi
[6 empty] | 1 | 2 | 3 | 4 | 5 | Unit
100 | µA | 10 | µA | 20 | Hz | 10 | 10 | 10 | 10 | 10 | µA | [6 empty] | 10 | µA | 0 | µA | 0.0001 | µA
100 | µA | 10 | µA | 50 | Hz | 10 | 10 | 10 | 10 | 10 | µA | [6 empty] | 10 | µA | 0 | µA | 0.0001 | µA
100 | µA | 10 | µA | 1 | kHz | 10 | 10 | 10 | 10 | 10 | µA | [6 empty] | 10 | µA | 0 | µA | 0.0001 | µA
100 | µA | 100 | µA | 20 | Hz | 100 | 100 | 100 | 100 | 100 | µA | [6 empty] | 100 | µA | 0 | µA | 0.0001 | µA
100 | µA | 100 | µA | 50 | Hz | 100 | 100 | 100 | 100 | 100 | µA | [6 empty] | 100 | µA | 0 | µA | 0.0001 | µA
100 | µA | 100 | µA | 1 | kHz | 100 | 100 | 100 | 100 | 100 | µA | [6 empty] | 100 | µA | 0 | µA | 0.0001 | µA
1 | mA | 0.1 | mA | 20 | Hz | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | mA | [6 empty] | 0.1 | mA | 0 | mA | 1e-06 | mA
1 | mA | 0.1 | mA | 50 | Hz | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | mA | [6 empty] | 0.1 | mA | 0 | mA | 1e-06 | mA
1 | mA | 0.1 | mA | 1 | kHz | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | mA | [6 empty] | 0.1 | mA | 0 | mA | 1e-06 | mA
1 | mA | 1 | mA | 20 | Hz | 1 | 1 | 1 | 1 | 1 | mA | [6 empty] | 1 | mA | 0 | mA | 1e-06 | mA
1 | mA | 1 | mA | 50 | Hz | 1 | 1 | 1 | 1 | 1 | mA | [6 empty] | 1 | mA | 0 | mA | 1e-06 | mA
1 | mA | 1 | mA | 1 | kHz | 1 | 1 | 1 | 1 | 1 | mA | [6 empty] | 1 | mA | 0 | mA | 1e-06 | mA
1 | mA | 1 | mA | 5 | kHz | 1 | 1 | 1 | 1 | 1 | mA | [6 empty] | 1 | mA | 0 | mA | 1e-06 | mA
1 | mA | 1 | mA | 10 | kHz | 1 | 1 | 1 | 1 | 1 | mA | [6 empty] | 1 | mA | 0 | mA | 1e-06 | mA
10 | mA | 1 | mA | 20 | Hz | 1 | 1 | 1 | 1 | 1 | mA | [6 empty] | 1 | mA | 0 | mA | 1e-05 | mA
10 | mA | 1 | mA | 50 | Hz | 1 | 1 | 1 | 1 | 1 | mA | [6 empty] | 1 | mA | 0 | mA | 1e-05 | mA
10 | mA | 1 | mA | 1 | kHz | 1 | 1 | 1 | 1 | 1 | mA | [6 empty] | 1 | mA | 0 | mA | 1e-05 | mA
10 | mA | 1 | mA | 5 | kHz | 1 | 1 | 1 | 1 | 1 | mA | [6 empty] | 1 | mA | 0 | mA | 1e-05 | mA
10 | mA | 1 | mA | 10 | kHz | 1 | 1 | 1 | 1 | 1 | mA | [6 empty] | 1 | mA | 0 | mA | 1e-05 | mA
10 | mA | 5 | mA | 20 | Hz | 5 | 5 | 5 | 5 | 5 | mA | [6 empty] | 5 | mA | 0 | mA | 1e-05 | mA
10 | mA | 5 | mA | 50 | Hz | 5 | 5 | 5 | 5 | 5 | mA | [6 empty] | 5 | mA | 0 | mA | 1e-05 | mA
10 | mA | 5 | mA | 1 | kHz | 5 | 5 | 5 | 5 | 5 | mA | [6 empty] | 5 | mA | 0 | mA | 1e-05 | mA
10 | mA | 5 | mA | 5 | kHz | 5 | 5 | 5 | 5 | 5 | mA | [6 empty] | 5 | mA | 0 | mA | 1e-05 | mA
10 | mA | 5 | mA | 10 | kHz | 5 | 5 | 5 | 5 | 5 | mA | [6 empty] | 5 | mA | 0 | mA | 1e-05 | mA
10 | mA | 10 | mA | 20 | Hz | 10 | 10 | 10 | 10 | 10 | mA | [6 empty] | 10 | mA | 0 | mA | 1e-05 | mA
10 | mA | 10 | mA | 50 | Hz | 10 | 10 | 10 | 10 | 10 | mA | [6 empty] | 10 | mA | 0 | mA | 1e-05 | mA
10 | mA | 10 | mA | 1 | kHz | 10 | 10 | 10 | 10 | 10 | mA | [6 empty] | 10 | mA | 0 | mA | 1e-05 | mA
10 | mA | 10 | mA | 5 | kHz | 10 | 10 | 10 | 10 | 10 | mA | [6 empty] | 10 | mA | 0 | mA | 1e-05 | mA
10 | mA | 10 | mA | 10 | kHz | 10 | 10 | 10 | 10 | 10 | mA | [6 empty] | 10 | mA | 0 | mA | 1e-05 | mA
100 | mA | 10 | mA | 20 | Hz | 10 | 10 | 10 | 10 | 10 | mA | [6 empty] | 10 | mA | 0 | mA | 0.0001 | mA
100 | mA | 10 | mA | 50 | Hz | 10 | 10 | 10 | 10 | 10 | mA | [6 empty] | 10 | mA | 0 | mA | 0.0001 | mA
100 | mA | 10 | mA | 1 | kHz | 10 | 10 | 10 | 10 | 10 | mA | [6 empty] | 10 | mA | 0 | mA | 0.0001 | mA
100 | mA | 10 | mA | 5 | kHz | 10 | 10 | 10 | 10 | 10 | mA | [6 empty] | 10 | mA | 0 | mA | 0.0001 | mA
100 | mA | 10 | mA | 10 | kHz | 10 | 10 | 10 | 10 | 10 | mA | [6 empty] | 10 | mA | 0 | mA | 0.0001 | mA
100 | mA | 100 | mA | 20 | Hz | 100 | 100 | 100 | 100 | 100 | mA | [6 empty] | 100 | mA | 0 | mA | 0.0001 | mA
100 | mA | 100 | mA | 50 | Hz | 100 | 100 | 100 | 100 | 100 | mA | [6 empty] | 100 | mA | 0 | mA | 0.0001 | mA
100 | mA | 100 | mA | 1 | kHz | 100 | 100 | 100 | 100 | 100 | mA | [6 empty] | 100 | mA | 0 | mA | 0.0001 | mA
100 | mA | 100 | mA | 5 | kHz | 100 | 100 | 100 | 100 | 100 | mA | [6 empty] | 100 | mA | 0 | mA | 0.0001 | mA
100 | mA | 100 | mA | 10 | kHz | 100 | 100 | 100 | 100 | 100 | mA | [6 empty] | 100 | mA | 0 | mA | 0.0001 | mA
1 | A | 0.1 | A | 20 | Hz | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | A | [6 empty] | 0.1 | A | 0 | A | 1e-06 | A
1 | A | 0.1 | A | 50 | Hz | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | A | [6 empty] | 0.1 | A | 0 | A | 1e-06 | A
1 | A | 0.1 | A | 1 | kHz | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | A | [6 empty] | 0.1 | A | 0 | A | 1e-06 | A
1 | A | 0.1 | A | 5 | kHz | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | A | [6 empty] | 0.1 | A | 0 | A | 1e-06 | A
1 | A | 0.1 | A | 10 | kHz | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | A | [6 empty] | 0.1 | A | 0 | A | 1e-06 | A
1 | A | 1 | A | 20 | Hz | 1 | 1 | 1 | 1 | 1 | A | [6 empty] | 1 | A | 0 | A | 1e-06 | A
1 | A | 1 | A | 50 | Hz | 1 | 1 | 1 | 1 | 1 | A | [6 empty] | 1 | A | 0 | A | 1e-06 | A
1 | A | 1 | A | 1 | kHz | 1 | 1 | 1 | 1 | 1 | A | [6 empty] | 1 | A | 0 | A | 1e-06 | A
1 | A | 1 | A | 5 | kHz | 1 | 1 | 1 | 1 | 1 | A | Rentang Pembacaan 3458 |  | Titik Ukur Tegangan 5730 |  | Frekuensi |  | 1 | A | 0 | A | 1e-06 | A
1 | A | 1 | A | 10 | kHz | 1 | 1 | 1 | 1 | 1 | A | [6 empty] | 1 | A | 0 | A | 1e-06 | A
30 | A | 3 | A | 20 | Hz | 3 | 3 | 3 | 3 | 3 | mV | 10 | mV | 0.3 | V | 20 | Hz | 3 | mV | 0 | mV | 1e-05 | mV
30 | A | 3 | A | 50 | Hz | 3 | 3 | 3 | 3 | 3 | mV | 10 | mV | 0.3 | V | 50 | Hz | 3 | mV | 0 | mV | 1e-05 | mV
30 | A | 3 | A | 1 | kHz | 3 | 3 | 3 | 3 | 3 | mV | 10 | mV | 0.3 | V | 1 | kHz | 3 | mV | 0 | mV | 1e-05 | mV
30 | A | 15 | A | 20 | Hz | 15 | 15 | 15 | 15 | 15 | mV | 100 | mV | 1.5 | V | 20 | Hz | 15 | mV | 0 | mV | 0.0001 | mV
30 | A | 15 | A | 50 | Hz | 15 | 15 | 15 | 15 | 15 | mV | 100 | mV | 1.5 | V | 50 | Hz | 15 | mV | 0 | mV | 0.0001 | mV
30 | A | 15 | A | 1 | kHz | 15 | 15 | 15 | 15 | 15 | mV | 100 | mV | 1.5 | V | 1 | kHz | 15 | mV | 0 | mV | 0.0001 | mV
30 | A | 20 | A | 20 | Hz | 20 | 20 | 20 | 20 | 20 | mV | 100 | mV | 2 | V | 20 | Hz | 20 | mV | 0 | mV | 0.0001 | mV
30 | A | 20 | A | 50 | Hz | 20 | 20 | 20 | 20 | 20 | mV | 100 | mV | 2 | V | 50 | Hz | 20 | mV | 0 | mV | 0.0001 | mV
30 | A | 20 | A | 1 | kHz | 20 | 20 | 20 | 20 | 20 | mV | 100 | mV | 2 | V | 1 | kHz | 20 | mV | 0 | mV | 0.0001 | mV
Resistansi 2W
Rentang UUC |  | Penunjukan Standar |  |  |  | Pembacaan UUC | [11 empty] | Rata-rata |  | Standar deviasi |  | Resolusi
[6 empty] | 1 | 2 | 3 | 4 | 5 | Unit
10 | Ω |  | 1 | Ω |  | 1 | 1 | 1 | 1 | 1 | Ω | [6 empty] | 1 | Ω | 0 | Ω | 1e-05 | Ω
10 | Ω |  | 10 | Ω |  | 10 | 10 | 10 | 10 | 10 | Ω | [6 empty] | 10 | Ω | 0 | Ω | 1e-05 | Ω
100 | Ω |  | 10 | Ω |  | 10 | 10 | 10 | 10 | 10 | Ω | [6 empty] | 10 | Ω | 0 | Ω | 1e-05 | Ω
100 | Ω |  | 100 | Ω |  | 100 | 100 | 100 | 100 | 100 | Ω | [6 empty] | 100 | Ω | 0 | Ω | 1e-05 | Ω
1 | kΩ |  | 0.1 | kΩ |  | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | kΩ | [6 empty] | 0.1 | kΩ | 0 | kΩ | 1e-07 | kΩ
1 | kΩ |  | 1 | kΩ |  | 1 | 1 | 1 | 1 | 1 | kΩ | [6 empty] | 1 | kΩ | 0 | kΩ | 1e-07 | kΩ
10 | kΩ |  | 1 | kΩ |  | 1 | 1 | 1 | 1 | 1 | kΩ | [6 empty] | 1 | kΩ | 0 | kΩ | 1e-06 | kΩ
10 | kΩ |  | 10 | kΩ |  | 10 | 10 | 10 | 10 | 10 | kΩ | [6 empty] | 10 | kΩ | 0 | kΩ | 1e-06 | kΩ
100 | kΩ |  | 10 | kΩ |  | 10 | 10 | 10 | 10 | 10 | kΩ | [6 empty] | 10 | kΩ | 0 | kΩ | 1e-05 | kΩ
100 | kΩ |  | 100 | kΩ |  | 100 | 100 | 100 | 100 | 100 | kΩ | [6 empty] | 100 | kΩ | 0 | kΩ | 1e-05 | kΩ
1 | MΩ |  | 0.1 | MΩ |  | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | MΩ | [6 empty] | 0.1 | MΩ | 0 | MΩ | 1e-07 | MΩ
1 | MΩ |  | 1 | MΩ |  | 1 | 1 | 1 | 1 | 1 | MΩ | [6 empty] | 1 | MΩ | 0 | MΩ | 1e-07 | MΩ
10 | MΩ |  | 1 | MΩ |  | 1 | 1 | 1 | 1 | 1 | MΩ | [6 empty] | 1 | MΩ | 0 | MΩ | 1e-06 | MΩ
10 | MΩ |  | 10 | MΩ |  | 10 | 10 | 10 | 10 | 10 | MΩ | [6 empty] | 10 | MΩ | 0 | MΩ | 1e-06 | MΩ
100 | MΩ |  | 10 | MΩ |  | 10 | 10 | 10 | 10 | 10 | MΩ | [6 empty] | 10 | MΩ | 0 | MΩ | 1e-05 | MΩ
100 | MΩ |  | 100 | MΩ |  | 100 | 100 | 100 | 100 | 100 | MΩ | [6 empty] | 100 | MΩ | 0 | MΩ | 1e-05 | MΩ
1 | GΩ |  | 0.1 | GΩ |  | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | GΩ | [6 empty] | 0.1 | GΩ | 0 | GΩ | 1e-07 | GΩ
1 | GΩ |  | 1 | GΩ |  | 1 | 1 | 1 | 1 | 1 | GΩ | [6 empty] | 1 | GΩ | 0 | GΩ | 1e-07 | GΩ

## Eval
Tegangan DC
Rentang UUC |  | Penunjukan Standar |  |  |  | Pembacaan alat |  | Koreksi Standar |  | Koreksi alat |  | ESDM UUC |  | RESOLUSI UUC |  | SERTIFIKAT STD |  | SPEK STD |  | TEMP STD |  | ROUNDING |  | EMF |  | Comb Unc |  | k | Exp Unc
100 | mV |  | 0 | mV |  | 0 | mV | -2.1073e-05 | mV | -2.1073e-05 | mV | 0 | mV | 2.8868e-06 | mV | 0.00058335 | mV | 0.00023094 | mV | 0.00011547 | mV | 2.8868e-05 | mV | 0.00018013 | mV | 0.00066352 | mV | 2 | 0.001327 | mV | 0.0001
100 | mV |  | 10 | mV |  | 10 | mV | -0.00011587 | mV | -0.00011587 | mV | 0 | mV | 2.8868e-06 | mV | 0.00058514 | mV | 0.00027424 | mV | 0.00012009 | mV | 2.8868e-05 | mV | 0.00018013 | mV | 0.00068213 | mV | 2 | 0.0013643 | mV | 0.0001
100 | mV |  | -10 | mV |  | -10 | mV | -2.3787e-05 | mV | -2.3787e-05 | mV | 0 | mV | 2.8868e-06 | mV | 0.00058542 | mV | 0.00027424 | mV | 0.00011085 | mV | 2.8868e-05 | mV | 0.00018013 | mV | 0.00068081 | mV | 2 | 0.0013616 | mV | 0.0001
100 | mV |  | 50 | mV |  | 50 | mV | -0.00049506 | mV | -0.00049506 | mV | 0 | mV | 2.8868e-06 | mV | 0.00060565 | mV | 0.00044745 | mV | 0.00013856 | mV | 2.8868e-05 | mV | 0.00018013 | mV | 0.00078709 | mV | 2 | 0.0015742 | mV | 0.0001
100 | mV |  | -50 | mV |  | -50 | mV | -7.1414e-05 | mV | -7.1414e-05 | mV | 0 | mV | 2.8868e-06 | mV | 0.00060565 | mV | 0.00044745 | mV | 9.2376e-05 | mV | 2.8868e-05 | mV | 0.00018013 | mV | 0.00078029 | mV | 2 | 0.0015606 | mV | 0.0001
100 | mV |  | 90 | mV |  | 90 | mV | -0.00087426 | mV | -0.00087426 | mV | 0 | mV | 2.8868e-06 | mV | 0.00062365 | mV | 0.00062065 | mV | 0.00015704 | mV | 2.8868e-05 | mV | 0.00018013 | mV | 0.00091219 | mV | 2 | 0.0018244 | mV | 0.0001
100 | mV |  | -90 | mV |  | -90 | mV | -0.00011904 | mV | -0.00011904 | mV | 0 | mV | 2.8868e-06 | mV | 0.0006237 | mV | 0.00062065 | mV | 7.3901e-05 | mV | 2.8868e-05 | mV | 0.00018013 | mV | 0.00090164 | mV | 2 | 0.0018033 | mV | 0.0001
100 | mV |  | 100 | mV |  | 100 | mV | -0.00096905 | mV | -0.00096905 | mV | 0 | mV | 2.8868e-06 | mV | 0.00062895 | mV | 0.00066395 | mV | 0.00016166 | mV | 2.8868e-05 | mV | 0.00018013 | mV | 0.00094649 | mV | 2 | 0.001893 | mV | 0.0001
100 | mV |  | -100 | mV |  | -100 | mV | -0.00013095 | mV | -0.00013095 | mV | 0 | mV | 2.8868e-06 | mV | 0.00062894 | mV | 0.00066395 | mV | 6.9282e-05 | mV | 2.8868e-05 | mV | 0.00018013 | mV | 0.00093514 | mV | 2 | 0.0018703 | mV | 0.0001
1 | V |  | 0.1 | V |  | 0.1 | V | -9.6905e-07 | V | -9.6905e-07 | V | 0 | V | 2.8868e-08 | V | 6.2895e-07 | V | 6.6395e-07 | V | 1.6166e-07 | V | 2.8868e-08 | V | 1.8013e-07 | V | 9.4692e-07 | V | 2 | 1.8938e-06 | V | 1e-07
1 | V |  | -0.1 | V |  | -0.1 | V | -1.3095e-07 | V | -1.3095e-07 | V | 0 | V | 2.8868e-08 | V | 6.2894e-07 | V | 6.6395e-07 | V | 6.9282e-08 | V | 2.8868e-08 | V | 1.8013e-07 | V | 9.3558e-07 | V | 2 | 1.8712e-06 | V | 1e-07
1 | V |  | 0.5 | V |  | 0.5 | V | -1.6602e-06 | V | -1.6602e-06 | V | 0 | V | 2.8868e-08 | V | 1.3971e-06 | V | 1.8475e-06 | V | 2.8868e-07 | V | 2.8868e-08 | V | 1.8013e-07 | V | 2.3415e-06 | V | 2 | 4.683e-06 | V | 1e-07
1 | V |  | -0.5 | V |  | -0.5 | V | 6.4034e-07 | V | 6.4034e-07 | V | 0 | V | 2.8868e-08 | V | 1.3982e-06 | V | 1.8475e-06 | V | -5.7735e-08 | V | 2.8868e-08 | V | 1.8013e-07 | V | 2.325e-06 | V | 2 | 4.65e-06 | V | 1e-07
1 | V |  | 0.9 | V |  | 0.9 | V | -2.0254e-06 | V | -2.0254e-06 | V | 0 | V | 2.8868e-08 | V | 1.9145e-06 | V | 3.0022e-06 | V | 4.2724e-07 | V | 2.8868e-08 | V | 1.8013e-07 | V | 3.591e-06 | V | 2 | 7.182e-06 | V | 1e-07
1 | V |  | -0.9 | V |  | -0.9 | V | 5.2655e-07 | V | 5.2655e-07 | V | 0 | V | 2.8868e-08 | V | 1.9154e-06 | V | 3.0022e-06 | V | -1.963e-07 | V | 2.8868e-08 | V | 1.8013e-07 | V | 3.5714e-06 | V | 2 | 7.1427e-06 | V | 1e-07
1 | V |  | 1 | V |  | 1 | V | -2.1167e-06 | V | -2.1167e-06 | V | 0 | V | 2.8868e-08 | V | 2.0546e-06 | V | 3.2909e-06 | V | 4.6188e-07 | V | 2.8868e-08 | V | 1.8013e-07 | V | 3.9114e-06 | V | 2 | 7.8228e-06 | V | 1e-07
1 | V |  | -1 | V |  | -1 | V | 4.981e-07 | V | 4.981e-07 | V | 0 | V | 2.8868e-08 | V | 2.0549e-06 | V | 3.2909e-06 | V | -2.3094e-07 | V | 2.8868e-08 | V | 1.8013e-07 | V | 3.891e-06 | V | 2 | 7.782e-06 | V | 1e-07
10 | V |  | 1 | V |  | 1 | V | -2.1167e-06 | V | -2.1167e-06 | V | 0 | V | 2.8868e-07 | V | 2.0546e-06 | V | 3.2909e-06 | V | 4.6188e-07 | V | 2.8868e-07 | V | 1.8013e-07 | V | 3.9324e-06 | V | 2 | 7.8648e-06 | V | 1e-06
10 | V |  | -1 | V |  | -1 | V | 4.981e-07 | V | 4.981e-07 | V | 0 | V | 2.8868e-07 | V | 2.0549e-06 | V | 3.2909e-06 | V | -2.3094e-07 | V | 2.8868e-07 | V | 1.8013e-07 | V | 3.9122e-06 | V | 2 | 7.8243e-06 | V | 1e-06
10 | V |  | 2 | V |  | 2 | V | -3.0297e-06 | V | -3.0297e-06 | V | 0 | V | 2.8868e-07 | V | 3.6056e-06 | V | 6.1776e-06 | V | 8.0829e-07 | V | 2.8868e-07 | V | 1.8013e-07 | V | 7.2122e-06 | V | 2 | 1.4424e-05 | V | 1e-06
10 | V |  | 3 | V |  | 3 | V | -1.0427e-05 | V | -1.0427e-05 | V | 0 | V | 2.8868e-07 | V | 1.1036e-05 | V | 7.5056e-06 | V | 7.5056e-07 | V | 2.8868e-07 | V | 1.8013e-07 | V | 1.3375e-05 | V | 2 | 2.6751e-05 | V | 1e-06
10 | V |  | 4 | V |  | 4 | V | -1.1942e-05 | V | -1.1942e-05 | V | 0 | V | 2.8868e-07 | V | 1.1036e-05 | V | 9.5263e-06 | V | 9.2376e-07 | V | 2.8868e-07 | V | 1.8013e-07 | V | 1.4615e-05 | V | 2 | 2.923e-05 | V | 1e-06
10 | V |  | 5 | V |  | 5 | V | -1.3458e-05 | V | -1.3458e-05 | V | 0 | V | 2.8868e-07 | V | 1.1036e-05 | V | 1.1547e-05 | V | 1.097e-06 | V | 2.8868e-07 | V | 1.8013e-07 | V | 1.6017e-05 | V | 2 | 3.2034e-05 | V | 1e-06
10 | V |  | -5 | V |  | -5 | V | 9.8766e-06 | V | 9.8766e-06 | V | 0 | V | 2.8868e-07 | V | 1.1113e-05 | V | 1.1547e-05 | V | -6.3509e-07 | V | 2.8868e-07 | V | 1.8013e-07 | V | 1.6045e-05 | V | 2 | 3.209e-05 | V | 1e-06
10 | V |  | 6 | V |  | 6 | V | -1.4973e-05 | V | -1.4973e-05 | V | 0 | V | 2.8868e-07 | V | 1.7885e-05 | V | 1.3568e-05 | V | 1.2702e-06 | V | 2.8868e-07 | V | 1.8013e-07 | V | 2.2489e-05 | V | 2 | 4.4978e-05 | V | 1e-06
10 | V |  | 7 | V |  | 7 | V | -1.6488e-05 | V | -1.6488e-05 | V | 0 | V | 2.8868e-07 | V | 1.7885e-05 | V | 1.5588e-05 | V | 1.4434e-06 | V | 2.8868e-07 | V | 1.8013e-07 | V | 2.3773e-05 | V | 2 | 4.7545e-05 | V | 1e-06
10 | V |  | 8 | V |  | 8 | V | -1.8003e-05 | V | -1.8003e-05 | V | 0 | V | 2.8868e-07 | V | 1.7885e-05 | V | 1.7609e-05 | V | 1.6166e-06 | V | 2.8868e-07 | V | 1.8013e-07 | V | 2.5155e-05 | V | 2 | 5.0309e-05 | V | 1e-06
10 | V |  | 9 | V |  | 9 | V | -1.9518e-05 | V | -1.9518e-05 | V | 0 | V | 2.8868e-07 | V | 1.7885e-05 | V | 1.963e-05 | V | 1.7898e-06 | V | 2.8868e-07 | V | 1.8013e-07 | V | 2.6619e-05 | V | 2 | 5.3239e-05 | V | 1e-06
10 | V |  | -9 | V |  | -9 | V | 8.0177e-06 | V | 8.0177e-06 | V | 0 | V | 2.8868e-07 | V | 1.7888e-05 | V | 1.963e-05 | V | -1.3279e-06 | V | 2.8868e-07 | V | 1.8013e-07 | V | 2.6595e-05 | V | 2 | 5.319e-05 | V | 1e-06
10 | V |  | 10 | V |  | 10 | V | -2.1033e-05 | V | -2.1033e-05 | V | 0 | V | 2.8868e-07 | V | 1.9367e-05 | V | 2.1651e-05 | V | 1.963e-06 | V | 2.8868e-07 | V | 1.8013e-07 | V | 2.9118e-05 | V | 2 | 5.8236e-05 | V | 1e-06
10 | V |  | -10 | V |  | -10 | V | 7.553e-06 | V | 7.553e-06 | V | 0 | V | 2.8868e-07 | V | 1.9372e-05 | V | 2.1651e-05 | V | -1.5011e-06 | V | 2.8868e-07 | V | 1.8013e-07 | V | 2.9094e-05 | V | 2 | 5.8188e-05 | V | 1e-06
100 | V |  | 10 | V |  | 10 | V | -2.1033e-05 | V | -2.1033e-05 | V | 0 | V | 2.8868e-06 | V | 1.9367e-05 | V | 2.1651e-05 | V | 1.963e-06 | V | 2.8868e-06 | V | 1.8013e-07 | V | 2.94e-05 | V | 2 | 5.88e-05 | V | 1e-05
100 | V |  | -10 | V |  | -10 | V | 7.553e-06 | V | 7.553e-06 | V | 0 | V | 2.8868e-06 | V | 1.9372e-05 | V | 2.1651e-05 | V | -1.5011e-06 | V | 2.8868e-06 | V | 1.8013e-07 | V | 2.9376e-05 | V | 2 | 5.8753e-05 | V | 1e-05
100 | V |  | 50 | V |  | 50 | V | -3.2271e-05 | V | -3.2271e-05 | V | 0 | V | 2.8868e-06 | V | 0.00015646 | V | 0.00016743 | V | 2.3094e-05 | V | 2.8868e-06 | V | 1.8013e-07 | V | 0.00023035 | V | 2 | 0.0004607 | V | 1e-05
100 | V |  | -50 | V |  | -50 | V | 4.7284e-05 | V | 4.7284e-05 | V | 0 | V | 2.8868e-06 | V | 0.00015644 | V | 0.00016743 | V | -1.1547e-05 | V | 2.8868e-06 | V | 1.8013e-07 | V | 0.00022947 | V | 2 | 0.00045894 | V | 1e-05
100 | V |  | 90 | V |  | 90 | V | -5.2675e-05 | V | -5.2675e-05 | V | 0 | V | 2.8868e-06 | V | 0.00024952 | V | 0.0002829 | V | 3.695e-05 | V | 2.8868e-06 | V | 1.8013e-07 | V | 0.00037905 | V | 2 | 0.00075809 | V | 1e-05
100 | V |  | -90 | V |  | -90 | V | 5.1501e-05 | V | 5.1501e-05 | V | 0 | V | 2.8868e-06 | V | 0.00024958 | V | 0.0002829 | V | -2.5403e-05 | V | 2.8868e-06 | V | 1.8013e-07 | V | 0.00037814 | V | 2 | 0.00075627 | V | 1e-05
100 | V |  | 100 | V |  | 100 | V | -5.7776e-05 | V | -5.7776e-05 | V | 0 | V | 2.8868e-06 | V | 0.00027332 | V | 0.00031177 | V | 4.0415e-05 | V | 2.8868e-06 | V | 1.8013e-07 | V | 0.0004166 | V | 2 | 0.00083319 | V | 1e-05
100 | V |  | -100 | V |  | -100 | V | 5.2556e-05 | V | 5.2556e-05 | V | 0 | V | 2.8868e-06 | V | 0.00027332 | V | 0.00031177 | V | -2.8868e-05 | V | 2.8868e-06 | V | 1.8013e-07 | V | 0.00041564 | V | 2 | 0.00083127 | V | 1e-05
1000 | V |  | 100 | V |  | 100 | V | -5.7776e-05 | V | -5.7776e-05 | V | 0 | V | 2.8868e-05 | V | 0.00027332 | V | 0.00031177 | V | 4.0415e-05 | V | 2.8868e-05 | V | 1.8013e-07 | V | 0.00041857 | V | 2 | 0.00083714 | V | 0.0001
1000 | V |  | -100 | V |  | -100 | V | 5.2556e-05 | V | 5.2556e-05 | V | 0 | V | 2.8868e-05 | V | 0.00027332 | V | 0.00031177 | V | -2.8868e-05 | V | 2.8868e-05 | V | 1.8013e-07 | V | 0.00041762 | V | 2 | 0.00083523 | V | 0.0001
1000 | V |  | 500 | V |  | 500 | V | -0.00063642 | V | -0.00063642 | V | 0 | V | 2.8868e-05 | V | 0.0015311 | V | 0.0014665 | V | 0.00030022 | V | 2.8868e-05 | V | 1.8013e-07 | V | 0.0021416 | V | 2 | 0.0042832 | V | 0.0001
1000 | V |  | -500 | V |  | -500 | V | 0.00013062 | V | 0.00013062 | V | 0 | V | 2.8868e-05 | V | 0.0015345 | V | 0.0014665 | V | -0.00027713 | V | 2.8868e-05 | V | 1.8013e-07 | V | 0.0021409 | V | 2 | 0.0042819 | V | 0.0001
1000 | V |  | 900 | V |  | 900 | V | -0.00076853 | V | -0.00076853 | V | 0 | V | 2.8868e-05 | V | 0.0027943 | V | 0.0026212 | V | 0.00053116 | V | 2.8868e-05 | V | 1.8013e-07 | V | 0.0038682 | V | 2 | 0.0077363 | V | 0.0001
1000 | V |  | -900 | V |  | -900 | V | 0.00033487 | V | 0.00033487 | V | 0 | V | 2.8868e-05 | V | 0.0027971 | V | 0.0026212 | V | -0.00050807 | V | 2.8868e-05 | V | 1.8013e-07 | V | 0.0038671 | V | 2 | 0.0077341 | V | 0.0001
1000 | V |  | 1000 | V |  | 1000 | V | -0.00080156 | V | -0.00080156 | V | 0 | V | 2.8868e-05 | V | 0.0027943 | V | 0.0029098 | V | 0.0005889 | V | 2.8868e-05 | V | 1.8013e-07 | V | 0.0040773 | V | 2 | 0.0081545 | V | 0.0001
1000 | V |  | -1000 | V |  | -1000 | V | 0.00038594 | V | 0.00038594 | V | 0 | V | 2.8868e-05 | V | 0.0027971 | V | 0.0029098 | V | -0.0005658 | V | 2.8868e-05 | V | 1.8013e-07 | V | 0.0040759 | V | 2 | 0.0081518 | V | 0.0001
Arus DC
Rentang UUC |  | Penunjukan Standar |  |  |  | Pembacaan alat |  | Koreksi Standar |  | Koreksi alat |  | ESDM UUC |  | RESOLUSI UUC |  | SERTIFIKAT STD |  | SPEK STD |  | TEMP STD |  | ROUNDING |  |  |  | Comb Unc |  | k | Exp Unc
100 | µA |  | 0 | µA |  | 0.001514 | µA | 0.0040992 | µA | 0.0025852 | µA | 0.0003812 | µA | 2.8868e-06 | µA | 0.0002336 | µA | 0.0034641 | µA | 0.00046188 | µA | 2.8868e-05 | µA |  |  | 0.0035234 | µA | 2 | 0.0070467 | µA | 0.0001
100 | µA |  | 10 | µA |  | 9.9984 | µA | 0.003967 | µA | 0.005581 | µA | 0.0004153 | µA | 2.8868e-06 | µA | 0.00027038 | µA | 0.003695 | µA | 0.00047343 | µA | 2.8868e-05 | µA |  |  | 0.0037582 | µA | 2 | 0.0075164 | µA | 0.0001
100 | µA |  | -10 | µA |  | -10.0013 | µA | 0.0043092 | µA | 0.0055712 | µA | 0.00031798 | µA | 2.8868e-06 | µA | 0.0002706 | µA | 0.003695 | µA | 0.00045033 | µA | 2.8868e-05 | µA |  |  | 0.0037458 | µA | 2 | 0.0074917 | µA | 0.0001
100 | µA |  | 50 | µA |  | 49.9976 | µA | 0.0034382 | µA | 0.0058202 | µA | 0.0005979 | µA | 2.8868e-06 | µA | 0.00044537 | µA | 0.0046188 | µA | 0.00051962 | µA | 2.8868e-05 | µA |  |  | 0.0047074 | µA | 2 | 0.0094149 | µA | 0.0001
100 | µA |  | -50 | µA |  | -50.0005 | µA | 0.0047642 | µA | 0.0052182 | µA | 0.00012331 | µA | 2.8868e-06 | µA | 0.00044543 | µA | 0.0046188 | µA | 0.00040415 | µA | 2.8868e-05 | µA |  |  | 0.0046595 | µA | 2 | 0.009319 | µA | 0.0001
100 | µA |  | 90 | µA |  | 89.9968 | µA | 0.0029094 | µA | 0.0060974 | µA | 0.00079851 | µA | 2.8868e-06 | µA | 0.00058971 | µA | 0.0055426 | µA | 0.0005658 | µA | 0.00028868 | µA |  |  | 0.0056665 | µA | 2 | 0.011333 | µA | 0.001
100 | µA |  | -90 | µA |  | -89.9998 | µA | 0.0052193 | µA | 0.0049793 | µA | 8.0685e-05 | µA | 2.8868e-06 | µA | 0.00058983 | µA | 0.0055426 | µA | 0.00035796 | µA | 0.00028868 | µA |  |  | 0.0055934 | µA | 2 | 0.011187 | µA | 0.001
100 | µA |  | 100 | µA |  | 99.9965 | µA | 0.0027772 | µA | 0.0062492 | µA | 0.00086928 | µA | 2.8868e-06 | µA | 0.00062632 | µA | 0.0057735 | µA | 0.00057735 | µA | 0.00028868 | µA |  |  | 0.0059074 | µA | 2 | 0.011815 | µA | 0.001
100 | µA |  | -100 | µA |  | -99.9996 | µA | 0.0053331 | µA | 0.0049611 | µA | 0.00011025 | µA | 2.8868e-06 | µA | 0.00062636 | µA | 0.0057735 | µA | 0.00034641 | µA | 0.00028868 | µA |  |  | 0.0058259 | µA | 2 | 0.011652 | µA | 0.001
1 | mA |  | 0.1 | mA |  | 0.099992 | mA | 2.7772e-06 | mA | 1.0597e-05 | mA | 1.9744e-06 | mA | 2.8868e-08 | mA | 6.2632e-07 | mA | 5.7735e-06 | mA | 5.7735e-07 | mA | 2.8868e-07 | mA |  |  | 6.1678e-06 | mA | 2 | 1.2336e-05 | mA | 1e-06
1 | mA |  | -0.1 | mA |  | -0.10002 | mA | 5.3331e-06 | mA | 2.3453e-05 | mA | 4.5636e-06 | mA | 2.8868e-08 | mA | 6.2636e-07 | mA | 5.7735e-06 | mA | 3.4641e-07 | mA | 2.8868e-07 | mA |  |  | 7.3997e-06 | mA | 2 | 1.4799e-05 | mA | 1e-06
1 | mA |  | 0.5 | mA |  | 0.50002 | mA | 5.8806e-06 | mA | -1.4319e-05 | mA | 5.0692e-06 | mA | 2.8868e-08 | mA | 4.4538e-06 | mA | 1.4145e-05 | mA | 2.8868e-06 | mA | 2.8868e-07 | mA |  |  | 1.5938e-05 | mA | 2 | 3.1877e-05 | mA | 1e-06
1 | mA |  | -0.5 | mA |  | -0.50004 | mA | 6.0705e-06 | mA | 4.8191e-05 | mA | 1.0538e-05 | mA | 2.8868e-08 | mA | 4.4531e-06 | mA | 1.4145e-05 | mA | 1.7321e-06 | mA | 2.8868e-07 | mA |  |  | 1.8277e-05 | mA | 2 | 3.6554e-05 | mA | 1e-06
1 | mA |  | 0.9 | mA |  | 0.90004 | mA | 5.9915e-06 | mA | -3.8929e-05 | mA | 1.1236e-05 | mA | 2.8868e-08 | mA | 5.8977e-06 | mA | 2.2228e-05 | mA | 3.3486e-06 | mA | 2.8868e-07 | mA |  |  | 2.5815e-05 | mA | 2 | 5.1629e-05 | mA | 1e-06
1 | mA |  | -0.9 | mA |  | -0.90007 | mA | 5.8566e-06 | mA | 7.1417e-05 | mA | 1.6397e-05 | mA | 2.8868e-08 | mA | 5.8971e-06 | mA | 2.2228e-05 | mA | 1.2702e-06 | mA | 2.8868e-07 | mA |  |  | 2.8274e-05 | mA | 2 | 5.6548e-05 | mA | 1e-06
1 | mA |  | 1 | mA |  | 1.0001 | mA | 6.0192e-06 | mA | -4.4761e-05 | mA | 1.2712e-05 | mA | 2.8868e-08 | mA | 6.2629e-06 | mA | 2.4249e-05 | mA | 3.4641e-06 | mA | 2.8868e-07 | mA |  |  | 2.83e-05 | mA | 2 | 5.66e-05 | mA | 1e-06
1 | mA |  | -1 | mA |  | -1.0001 | mA | 5.8031e-06 | mA | 7.7983e-05 | mA | 1.8051e-05 | mA | 2.8868e-08 | mA | 6.2632e-06 | mA | 2.4249e-05 | mA | 1.1547e-06 | mA | 2.8868e-07 | mA |  |  | 3.0895e-05 | mA | 2 | 6.179e-05 | mA | 1e-06
10 | mA |  | 1 | mA |  | 0.99988 | mA | 6.0192e-06 | mA | 0.00012542 | mA | 3.0336e-05 | mA | 2.8868e-07 | mA | 6.2629e-06 | mA | 2.4249e-05 | mA | 3.4641e-06 | mA | 2.8868e-07 | mA |  |  | 3.9492e-05 | mA | 2 | 7.8985e-05 | mA | 1e-06
10 | mA |  | -1 | mA |  | -1.0001 | mA | 5.8031e-06 | mA | 0.000153 | mA | 3.6857e-05 | mA | 2.8868e-07 | mA | 6.2632e-06 | mA | 2.4249e-05 | mA | 1.1547e-06 | mA | 2.8868e-07 | mA |  |  | 4.4578e-05 | mA | 2 | 8.9155e-05 | mA | 1e-06
10 | mA |  | 2 | mA |  | 1.9999 | mA | 6.2963e-06 | mA | 0.0001119 | mA | 2.6931e-05 | mA | 2.8868e-07 | mA | 9.9002e-06 | mA | 4.4456e-05 | mA | 4.6188e-06 | mA | 2.8868e-07 | mA |  |  | 5.3114e-05 | mA | 2 | 0.00010623 | mA | 1e-06
10 | mA |  | 3 | mA |  | 2.9999 | mA | 3.8686e-05 | mA | 8.8886e-05 | mA | 1.3021e-05 | mA | 2.8868e-07 | mA | 4.759e-05 | mA | 8.3716e-05 | mA | 2.6558e-05 | mA | 2.8868e-06 | mA |  |  | 0.00010078 | mA | 2 | 0.00020156 | mA | 1e-05
10 | mA |  | 4 | mA |  | 4 | mA | 4.4743e-05 | mA | 7.3743e-05 | mA | 8.0808e-06 | mA | 2.8868e-07 | mA | 4.759e-05 | mA | 0.00010392 | mA | 2.7713e-05 | mA | 2.8868e-06 | mA |  |  | 0.00011793 | mA | 2 | 0.00023585 | mA | 1e-05
10 | mA |  | 5 | mA |  | 5 | mA | 5.08e-05 | mA | 5.46e-05 | mA | 5.9195e-06 | mA | 2.8868e-07 | mA | 4.759e-05 | mA | 0.00012413 | mA | 2.8868e-05 | mA | 2.8868e-06 | mA |  |  | 0.0001362 | mA | 2 | 0.0002724 | mA | 1e-05
10 | mA |  | -5 | mA |  | -5.0002 | mA | -7.4223e-06 | mA | 0.00020258 | mA | 5.2615e-05 | mA | 2.8868e-07 | mA | 4.759e-05 | mA | 0.00012413 | mA | 1.7321e-05 | mA | 2.8868e-06 | mA |  |  | 0.00014405 | mA | 2 | 0.0002881 | mA | 1e-05
10 | mA |  | 6 | mA |  | 6 | mA | 5.6858e-05 | mA | 4.4258e-05 | mA | 5.0853e-06 | mA | 2.8868e-07 | mA | 6.7588e-05 | mA | 0.00014434 | mA | 3.0022e-05 | mA | 2.8868e-06 | mA |  |  | 0.00016229 | mA | 2 | 0.00032457 | mA | 1e-05
10 | mA |  | 7 | mA |  | 7 | mA | 6.2915e-05 | mA | 3.5715e-05 | mA | 9.5572e-06 | mA | 2.8868e-07 | mA | 6.7588e-05 | mA | 0.00016454 | mA | 3.1177e-05 | mA | 2.8868e-06 | mA |  |  | 0.00018087 | mA | 2 | 0.00036174 | mA | 1e-05
10 | mA |  | 8 | mA |  | 8 | mA | 6.8972e-05 | mA | 2.0772e-05 | mA | 1.271e-05 | mA | 2.8868e-07 | mA | 6.7588e-05 | mA | 0.00018475 | mA | 3.2332e-05 | mA | 2.8868e-06 | mA |  |  | 0.00019979 | mA | 2 | 0.00039958 | mA | 1e-05
10 | mA |  | 9 | mA |  | 9.0001 | mA | 7.5029e-05 | mA | 8.629e-06 | mA | 1.7137e-05 | mA | 2.8868e-07 | mA | 6.7588e-05 | mA | 0.00020496 | mA | 3.3486e-05 | mA | 2.8868e-06 | mA |  |  | 0.00021909 | mA | 2 | 0.00043818 | mA | 1e-05
10 | mA |  | -9 | mA |  | -9.0003 | mA | -3.892e-05 | mA | 0.00025028 | mA | 7.2455e-05 | mA | 2.8868e-07 | mA | 6.7554e-05 | mA | 0.00020496 | mA | 1.2702e-05 | mA | 2.8868e-06 | mA |  |  | 0.00022802 | mA | 2 | 0.00045603 | mA | 1e-05
10 | mA |  | 10 | mA |  | 10.0001 | mA | 8.1086e-05 | mA | -1.0914e-05 | mA | 2.33e-05 | mA | 2.8868e-07 | mA | 7.2233e-05 | mA | 0.00022517 | mA | 3.4641e-05 | mA | 2.8868e-06 | mA |  |  | 0.00024014 | mA | 2 | 0.00048029 | mA | 1e-05
10 | mA |  | -10 | mA |  | -10.0003 | mA | -4.6795e-05 | mA | 0.00027041 | mA | 7.9615e-05 | mA | 2.8868e-07 | mA | 7.2236e-05 | mA | 0.00022517 | mA | 1.1547e-05 | mA | 2.8868e-06 | mA |  |  | 0.0002498 | mA | 2 | 0.00049959 | mA | 1e-05
100 | mA |  | 10 | mA |  | 10.0002 | mA | 8.1086e-05 | mA | -7.6914e-05 | mA | 6.0283e-05 | mA | 2.8868e-06 | mA | 7.2233e-05 | mA | 0.00022517 | mA | 3.4641e-05 | mA | 2.8868e-06 | mA |  |  | 0.00024651 | mA | 2 | 0.00049302 | mA | 1e-05
100 | mA |  | -10 | mA |  | -10.0027 | mA | -4.6795e-05 | mA | 0.0026632 | mA | 0.00067918 | mA | 2.8868e-06 | mA | 7.2236e-05 | mA | 0.00022517 | mA | 1.1547e-05 | mA | 2.8868e-06 | mA |  |  | 0.00071928 | mA | 2 | 0.0014386 | mA | 1e-05
100 | mA |  | 50 | mA |  | 50.0063 | mA | 0.0010634 | mA | -0.0052086 | mA | 0.0015683 | mA | 2.8868e-06 | mA | 0.0014642 | mA | 0.0012994 | mA | 0.00028868 | mA | 2.8868e-05 | mA |  |  | 0.0025251 | mA | 2 | 0.0050501 | mA | 0.0001
100 | mA |  | -50 | mA |  | -50.0083 | mA | -8.1135e-06 | mA | 0.0082979 | mA | 0.0020771 | mA | 2.8868e-06 | mA | 0.0014642 | mA | 0.0012994 | mA | 0.00017321 | mA | 2.8868e-05 | mA |  |  | 0.0028597 | mA | 2 | 0.0057193 | mA | 0.0001
100 | mA |  | 90 | mA |  | 90.0123 | mA | 0.0015814 | mA | -0.010703 | mA | 0.0030712 | mA | 2.8868e-06 | mA | 0.0022694 | mA | 0.0023387 | mA | 0.00033486 | mA | 2.8868e-05 | mA |  |  | 0.0044905 | mA | 2 | 0.008981 | mA | 0.0001
100 | mA |  | -90 | mA |  | -90.0145 | mA | -1.9198e-05 | mA | 0.014525 | mA | 0.0036362 | mA | 2.8868e-06 | mA | 0.0022694 | mA | 0.0023387 | mA | 0.00012702 | mA | 2.8868e-05 | mA |  |  | 0.0048845 | mA | 2 | 0.0097691 | mA | 0.0001
100 | mA |  | 100 | mA |  | 100.0138 | mA | 0.0017109 | mA | -0.012105 | mA | 0.0034541 | mA | 2.8868e-06 | mA | 0.0024712 | mA | 0.0025985 | mA | 0.00034641 | mA | 2.8868e-05 | mA |  |  | 0.0049911 | mA | 2 | 0.0099822 | mA | 0.0001
100 | mA |  | -100 | mA |  | -100.0161 | mA | -2.1969e-05 | mA | 0.01612 | mA | 0.0040358 | mA | 2.8868e-06 | mA | 0.0024713 | mA | 0.0025985 | mA | 0.00011547 | mA | 2.8868e-05 | mA |  |  | 0.0054001 | mA | 2 | 0.0108 | mA | 0.0001
1 | A |  | 0.1 | A |  | 0.099984 | A | 1.7109e-06 | A | 1.7751e-05 | A | 4.0394e-06 | A | 2.8868e-08 | A | 2.4712e-06 | A | 2.5985e-06 | A | 3.4641e-07 | A | 2.8868e-08 | A |  |  | 5.4127e-06 | A | 2 | 1.0825e-05 | A | 1e-07
1 | A |  | -0.1 | A |  | -0.10001 | A | -2.1969e-08 | A | 1.4198e-05 | A | 3.5829e-06 | A | 2.8868e-08 | A | 2.4713e-06 | A | 2.5985e-06 | A | 1.1547e-07 | A | 2.8868e-08 | A |  |  | 5.0707e-06 | A | 2 | 1.0141e-05 | A | 1e-07
1 | A |  | 0.5 | A |  | 0.49999 | A | -1.0956e-05 | A | -1.6555e-06 | A | 7.7067e-06 | A | 2.8868e-08 | A | 5.8704e-05 | A | 2.3101e-05 | A | 3.4641e-06 | A | 2.8868e-06 | A |  |  | 6.3714e-05 | A | 2 | 0.00012743 | A | 1e-05
1 | A |  | -0.5 | A |  | -0.50001 | A | 2.9587e-05 | A | 4.3267e-05 | A | 4.9191e-06 | A | 2.8868e-08 | A | 5.8703e-05 | A | 2.3101e-05 | A | 2.3094e-06 | A | 2.8868e-06 | A |  |  | 6.3384e-05 | A | 2 | 0.00012677 | A | 1e-05
1 | A |  | 0.9 | A |  | 0.89997 | A | -2.6041e-05 | A | 3.1587e-06 | A | 8.0143e-06 | A | 2.8868e-08 | A | 9.8044e-05 | A | 4.1576e-05 | A | 3.926e-06 | A | 2.8868e-06 | A |  |  | 0.00010691 | A | 2 | 0.00021382 | A | 1e-05
1 | A |  | -0.9 | A |  | -0.9 | A | 4.3296e-05 | A | 4.2176e-05 | A | 2.1544e-06 | A | 2.8868e-08 | A | 9.8044e-05 | A | 4.1576e-05 | A | 1.8475e-06 | A | 2.8868e-06 | A |  |  | 0.00010657 | A | 2 | 0.00021314 | A | 1e-05
1 | A |  | 1 | A |  | 0.99997 | A | -2.9813e-05 | A | -1.4528e-06 | A | 7.3921e-06 | A | 2.8868e-08 | A | 0.00010789 | A | 4.6195e-05 | A | 4.0415e-06 | A | 2.8868e-06 | A |  |  | 0.0001177 | A | 2 | 0.0002354 | A | 1e-05
1 | A |  | -1 | A |  | -0.99998 | A | 4.6723e-05 | A | 3.0943e-05 | A | 7.4329e-06 | A | 2.8868e-08 | A | 0.00010789 | A | 4.6195e-05 | A | 1.7321e-06 | A | 2.8868e-06 | A |  |  | 0.00011765 | A | 2 | 0.00023529 | A | 1e-05
Tegangan AC
Rentang UUC |  | Penunjukan Standar |  |  |  | Pembacaan alat |  | Koreksi Standar |  | Koreksi alat |  | ESDM UUC |  | RESOLUSI UUC |  | SERTIFIKAT STD |  | SPEK STD |  | TEMP STD |  | ROUNDING |  |  |  | Comb Unc |  | k | Exp Unc
10 | mV | 10 | mV | 20 | Hz | 10 | mV | -0.00016155 | mV | -0.00016155 | mV | 0 | mV | 2.8868e-06 | mV | -8.1861e-05 | mV | 0.002829 | mV | 0.00034641 | mV | 2.8868e-05 | mV |  |  | 0.0028515 | mV | 2 | 0.0057029 | mV | 0.0001
10 | mV | 10 | mV | 50 | Hz | 10 | mV | -0.00041446 | mV | -0.00041446 | mV | 0 | mV | 2.8868e-06 | mV | -0.00020817 | mV | 0.0027713 | mV | 0.00034641 | mV | 2.8868e-05 | mV |  |  | 0.0028007 | mV | 2 | 0.0056015 | mV | 0.0001
10 | mV | 10 | mV | 1 | kHz | 10 | mV | -0.00041446 | mV | -0.00041446 | mV | 0 | mV | 2.8868e-06 | mV | -0.00020817 | mV | 0.0027713 | mV | 0.00034641 | mV | 2.8868e-05 | mV |  |  | 0.0028007 | mV | 2 | 0.0056015 | mV | 0.0001
10 | mV | 10 | mV | 10 | kHz | 10 | mV | -0.0009925 | mV | -0.0009925 | mV | 0 | mV | 2.8868e-06 | mV | -0.00049766 | mV | 0.0027713 | mV | 0.00034641 | mV | 2.8868e-05 | mV |  |  | 0.002837 | mV | 2 | 0.005674 | mV | 0.0001
10 | mV | 10 | mV | 20 | kHz | 10 | mV | -0.0052128 | mV | -0.0052128 | mV | 0 | mV | 2.8868e-06 | mV | -0.002607 | mV | 0.0034641 | mV | 0.00057735 | mV | 2.8868e-05 | mV |  |  | 0.0043738 | mV | 2 | 0.0087477 | mV | 0.0001
100 | mV | 20 | mV | 50 | kHz | 20 | mV | -0.012276 | mV | -0.012276 | mV | 0 | mV | 2.8868e-05 | mV | -0.002607 | mV | 0.0086603 | mV | 0.00057735 | mV | 0.00028868 | mV |  |  | 0.0090672 | mV | 2 | 0.018134 | mV | 0.001
100 | mV | 30 | mV | 100 | kHz | 30 | mV | -0.026265 | mV | -0.026265 | mV | 0 | mV | 2.8868e-05 | mV | -0.032852 | mV | 0.022892 | mV | 0.020092 | mV | 0.00028868 | mV |  |  | 0.0448 | mV | 2 | 0.089601 | mV | 0.001
100 | mV | 50 | mV | 40 | Hz | 50 | mV | 0.00033009 | mV | 0.00033009 | mV | 0 | mV | 2.8868e-05 | mV | 0.00029922 | mV | 0.0056869 | mV | 0.0012702 | mV | 0.00028868 | mV |  |  | 0.0058419 | mV | 2 | 0.011684 | mV | 0.001
100 | mV | 100 | mV | 80 | Hz | 100 | mV | 8.6035e-06 | mV | 8.6035e-06 | mV | 0 | mV | 2.8868e-05 | mV | 4.0471e-06 | mV | 0.0073323 | mV | 0.0013856 | mV | 0.00028868 | mV |  |  | 0.0074678 | mV | 2 | 0.014936 | mV | 0.001
100 | mV | 100 | mV | 20 | Hz | 100 | mV | 0.00059783 | mV | 0.00059783 | mV | 0 | mV | 2.8868e-05 | mV | 0.0082478 | mV | 0.0092376 | mV | 0.0013856 | mV | 0.00028868 | mV |  |  | 0.012465 | mV | 2 | 0.024929 | mV | 0.001
100 | mV | 100 | mV | 50 | Hz | 100 | mV | 8.6035e-06 | mV | 8.6035e-06 | mV | 0 | mV | 2.8868e-05 | mV | 0.0071189 | mV | 0.0073323 | mV | 0.0013856 | mV | 0.00028868 | mV |  |  | 0.010317 | mV | 2 | 0.020635 | mV | 0.001
100 | mV | 100 | mV | 1 | kHz | 100 | mV | 8.6035e-06 | mV | 8.6035e-06 | mV | 0 | mV | 2.8868e-05 | mV | 0.0071189 | mV | 0.0073323 | mV | 0.0013856 | mV | 0.00028868 | mV |  |  | 0.010317 | mV | 2 | 0.020635 | mV | 0.001
100 | mV | 100 | mV | 10 | kHz | 100 | mV | -0.012419 | mV | -0.012419 | mV | 0 | mV | 2.8868e-05 | mV | 0.022104 | mV | 0.0073323 | mV | 0.0013856 | mV | 0.00028868 | mV |  |  | 0.023331 | mV | 2 | 0.046663 | mV | 0.001
100 | mV | 100 | mV | 20 | kHz | 100 | mV | 0.10445 | mV | 0.10445 | mV | 0 | mV | 2.8868e-05 | mV | 0.052207 | mV | 0.01097 | mV | 0.0040415 | mV | 0.0028868 | mV |  |  | 0.053578 | mV | 2 | 0.10716 | mV | 0.01
100 | mV | 100 | mV | 50 | kHz | 100 | mV | 0.10445 | mV | 0.10445 | mV | 0 | mV | 2.8868e-05 | mV | 0.052207 | mV | 0.027713 | mV | 0.0063509 | mV | 0.0028868 | mV |  |  | 0.059517 | mV | 2 | 0.11903 | mV | 0.01
100 | mV | 100 | mV | 100 | kHz | 100 | mV | -0.065749 | mV | -0.065749 | mV | 0 | mV | 2.8868e-05 | mV | 0.052373 | mV | 0.049363 | mV | 0.026558 | mV | 0.0028868 | mV |  |  | 0.076768 | mV | 2 | 0.15354 | mV | 0.01
1 | V | 0.1 | V | 20 | Hz | 0.1 | V | 5.9783e-07 | V | 5.9783e-07 | V | 0 | V | 2.8868e-07 | V | 8.2478e-06 | V | 9.2376e-06 | V | 1.3856e-06 | V | 2.8868e-07 | V |  |  | 1.2468e-05 | V | 2 | 2.4936e-05 | V | 1e-06
1 | V | 0.1 | V | 50 | Hz | 0.1 | V | 8.6035e-09 | V | 8.6035e-09 | V | 0 | V | 2.8868e-07 | V | 7.1189e-06 | V | 7.3323e-06 | V | 1.3856e-06 | V | 2.8868e-07 | V |  |  | 1.0321e-05 | V | 2 | 2.0643e-05 | V | 1e-06
1 | V | 0.1 | V | 1 | kHz | 0.1 | V | 8.6035e-09 | V | 8.6035e-09 | V | 0 | V | 2.8868e-07 | V | 7.1189e-06 | V | 7.3323e-06 | V | 1.3856e-06 | V | 2.8868e-07 | V |  |  | 1.0321e-05 | V | 2 | 2.0643e-05 | V | 1e-06
1 | V | 0.1 | V | 10 | kHz | 0.1 | V | -1.2419e-05 | V | -1.2419e-05 | V | 0 | V | 2.8868e-07 | V | 2.2104e-05 | V | 7.3323e-06 | V | 1.3856e-06 | V | 2.8868e-07 | V |  |  | 2.3333e-05 | V | 2 | 4.6666e-05 | V | 1e-06
1 | V | 0.1 | V | 20 | kHz | 0.1 | V | 0.00010445 | V | 0.00010445 | V | 0 | V | 2.8868e-07 | V | 5.2207e-05 | V | 1.097e-05 | V | 4.0415e-06 | V | 2.8868e-06 | V |  |  | 5.3578e-05 | V | 2 | 0.00010716 | V | 1e-05
1 | V | 0.1 | V | 50 | kHz | 0.1 | V | 0.00010445 | V | 0.00010445 | V | 0 | V | 2.8868e-07 | V | 5.2207e-05 | V | 2.7713e-05 | V | 6.3509e-06 | V | 2.8868e-06 | V |  |  | 5.9517e-05 | V | 2 | 0.00011903 | V | 1e-05
1 | V | 0.1 | V | 100 | kHz | 0.1 | V | -6.5749e-05 | V | -6.5749e-05 | V | 0 | V | 2.8868e-07 | V | 5.2373e-05 | V | 4.9363e-05 | V | 2.6558e-05 | V | 2.8868e-06 | V |  |  | 7.6769e-05 | V | 2 | 0.00015354 | V | 1e-05
1 | V | 0.3 | V | 80 | Hz | 0.3 | V | 2.5641e-05 | V | 2.5641e-05 | V | 0 | V | 2.8868e-07 | V | 4.6978e-05 | V | 1.1893e-05 | V | 4.0415e-06 | V | 2.8868e-07 | V |  |  | 4.863e-05 | V | 2 | 9.726e-05 | V | 1e-06
1 | V | 0.4 | V | 100 | Hz | 0.4 | V | 3.0924e-05 | V | 3.0924e-05 | V | 0 | V | 2.8868e-07 | V | 4.6978e-05 | V | 1.4318e-05 | V | 4.6188e-06 | V | 2.8868e-07 | V |  |  | 4.933e-05 | V | 2 | 9.866e-05 | V | 1e-06
1 | V | 0.5 | V | 160 | Hz | 0.5 | V | 3.6208e-05 | V | 3.6208e-05 | V | 0 | V | 2.8868e-07 | V | 4.6978e-05 | V | 1.6743e-05 | V | 5.1962e-06 | V | 2.8868e-06 | V |  |  | 5.0226e-05 | V | 2 | 0.00010045 | V | 1e-05
1 | V | 0.5 | V | 200 | Hz | 0.5 | V | 3.6208e-05 | V | 3.6208e-05 | V | 0 | V | 2.8868e-07 | V | 4.6978e-05 | V | 1.6743e-05 | V | 5.1962e-06 | V | 2.8868e-06 | V |  |  | 5.0226e-05 | V | 2 | 0.00010045 | V | 1e-05
1 | V | 0.6 | V | 400 | Hz | 0.6 | V | 4.1492e-05 | V | 4.1492e-05 | V | 0 | V | 2.8868e-07 | V | 4.6978e-05 | V | 1.9168e-05 | V | 5.7735e-06 | V | 2.8868e-06 | V |  |  | 5.1148e-05 | V | 2 | 0.0001023 | V | 1e-05
1 | V | 0.6 | V | 500 | Hz | 0.6 | V | 4.1492e-05 | V | 4.1492e-05 | V | 0 | V | 2.8868e-07 | V | 4.6978e-05 | V | 1.9168e-05 | V | 5.7735e-06 | V | 2.8868e-06 | V |  |  | 5.1148e-05 | V | 2 | 0.0001023 | V | 1e-05
1 | V | 0.8 | V | 1000 | Hz | 0.8 | V | 5.2059e-05 | V | 5.2059e-05 | V | 0 | V | 2.8868e-07 | V | 4.6978e-05 | V | 2.4018e-05 | V | 6.9282e-06 | V | 2.8868e-06 | V |  |  | 5.3294e-05 | V | 2 | 0.00010659 | V | 1e-05
1 | V | 0.8 | V | 1600 | Hz | 0.8 | V | 0.00010493 | V | 0.00010493 | V | 0 | V | 2.8868e-07 | V | 5.9246e-05 | V | 2.4018e-05 | V | 6.9282e-06 | V | 2.8868e-06 | V |  |  | 6.4369e-05 | V | 2 | 0.00012874 | V | 1e-05
1 | V | 1 | V | 20 | Hz | 1 | V | 7.1456e-05 | V | 7.1456e-05 | V | 0 | V | 2.8868e-07 | V | 5.9114e-05 | V | 6.0622e-05 | V | 2.3094e-05 | V | 2.8868e-06 | V |  |  | 8.7814e-05 | V | 2 | 0.00017563 | V | 1e-05
1 | V | 1 | V | 50 | Hz | 1 | V | 6.2626e-05 | V | 6.2626e-05 | V | 0 | V | 2.8868e-07 | V | 4.6978e-05 | V | 2.8868e-05 | V | 8.0829e-06 | V | 2.8868e-06 | V |  |  | 5.5803e-05 | V | 2 | 0.00011161 | V | 1e-05
1 | V | 1 | V | 1 | kHz | 1 | V | 6.2626e-05 | V | 6.2626e-05 | V | 0 | V | 2.8868e-07 | V | 4.6978e-05 | V | 2.8868e-05 | V | 8.0829e-06 | V | 2.8868e-06 | V |  |  | 5.5803e-05 | V | 2 | 0.00011161 | V | 1e-05
1 | V | 1 | V | 10 | kHz | 1 | V | -5.8781e-05 | V | -5.8781e-05 | V | 0 | V | 2.8868e-07 | V | 0.00013929 | V | 2.8868e-05 | V | 8.0829e-06 | V | 2.8868e-06 | V |  |  | 0.00014251 | V | 2 | 0.00028502 | V | 1e-05
1 | V | 1 | V | 20 | kHz | 1 | V | -0.00051695 | V | -0.00051695 | V | 0 | V | 2.8868e-07 | V | 0.00040677 | V | 4.4456e-05 | V | 2.1939e-05 | V | 2.8868e-06 | V |  |  | 0.00040979 | V | 2 | 0.00081959 | V | 1e-05
1 | V | 1 | V | 50 | kHz | 1 | V | -0.00051695 | V | -0.00051695 | V | 0 | V | 2.8868e-07 | V | 0.00040677 | V | 0.00016454 | V | 2.7713e-05 | V | 2.8868e-06 | V |  |  | 0.00043968 | V | 2 | 0.00087936 | V | 1e-05
1 | V | 1 | V | 100 | kHz | 1 | V | -0.00046656 | V | -0.00046656 | V | 0 | V | 2.8868e-07 | V | 0.00040686 | V | 0.00024018 | V | 0.0001097 | V | 2.8868e-06 | V |  |  | 0.00048504 | V | 2 | 0.00097007 | V | 1e-05
10 | V | 1 | V | 20 | Hz | 1 | V | 7.1456e-05 | V | 7.1456e-05 | V | 0 | V | 2.8868e-06 | V | 5.9114e-05 | V | 6.0622e-05 | V | 2.3094e-05 | V | 2.8868e-06 | V |  |  | 8.7861e-05 | V | 2 | 0.00017572 | V | 1e-05
10 | V | 1 | V | 50 | Hz | 1 | V | 6.2626e-05 | V | 6.2626e-05 | V | 0 | V | 2.8868e-06 | V | 4.6978e-05 | V | 2.8868e-05 | V | 8.0829e-06 | V | 2.8868e-06 | V |  |  | 5.5877e-05 | V | 2 | 0.00011175 | V | 1e-05
10 | V | 1 | V | 1 | kHz | 1 | V | 6.2626e-05 | V | 6.2626e-05 | V | 0 | V | 2.8868e-06 | V | 4.6978e-05 | V | 2.8868e-05 | V | 8.0829e-06 | V | 2.8868e-06 | V |  |  | 5.5877e-05 | V | 2 | 0.00011175 | V | 1e-05
10 | V | 1 | V | 2 | kHz | 1 | V | 0.00013003 | V | 0.00013003 | V | 0 | V | 2.8868e-06 | V | 5.9246e-05 | V | 2.8868e-05 | V | 8.0829e-06 | V | 2.8868e-06 | V |  |  | 6.6524e-05 | V | 2 | 0.00013305 | V | 1e-05
10 | V | 1 | V | 4 | kHz | 1 | V | 0.00013003 | V | 0.00013003 | V | 0 | V | 2.8868e-06 | V | 5.9246e-05 | V | 2.8868e-05 | V | 8.0829e-06 | V | 2.8868e-06 | V |  |  | 6.6524e-05 | V | 2 | 0.00013305 | V | 1e-05
10 | V | 1 | V | 5 | kHz | 1 | V | 0.00013003 | V | 0.00013003 | V | 0 | V | 2.8868e-06 | V | 5.9246e-05 | V | 2.8868e-05 | V | 8.0829e-06 | V | 2.8868e-06 | V |  |  | 6.6524e-05 | V | 2 | 0.00013305 | V | 1e-05
10 | V | 1 | V | 10 | kHz | 1 | V | -5.8781e-05 | V | -5.8781e-05 | V | 0 | V | 2.8868e-06 | V | 0.00013929 | V | 2.8868e-05 | V | 8.0829e-06 | V | 2.8868e-06 | V |  |  | 0.00014254 | V | 2 | 0.00028508 | V | 1e-05
10 | V | 1 | V | 20 | kHz | 1 | V | -0.00051695 | V | -0.00051695 | V | 0 | V | 2.8868e-06 | V | 0.00040677 | V | 4.4456e-05 | V | 2.1939e-05 | V | 2.8868e-06 | V |  |  | 0.0004098 | V | 2 | 0.00081961 | V | 1e-05
10 | V | 1 | V | 50 | kHz | 1 | V | -0.00051695 | V | -0.00051695 | V | 0 | V | 2.8868e-06 | V | 0.00040677 | V | 0.00016454 | V | 2.7713e-05 | V | 2.8868e-06 | V |  |  | 0.00043969 | V | 2 | 0.00087937 | V | 1e-05
10 | V | 1 | V | 100 | kHz | 1 | V | -0.00046656 | V | -0.00046656 | V | 0 | V | 2.8868e-06 | V | 0.00040686 | V | 0.00024018 | V | 0.0001097 | V | 2.8868e-06 | V |  |  | 0.00048504 | V | 2 | 0.00097009 | V | 1e-05
10 | V | 5 | V | 20 | Hz | 5 | V | 0.00038142 | V | 0.00038142 | V | 0 | V | 2.8868e-06 | V | 0.00058666 | V | 0.00034641 | V | 0.00012124 | V | 2.8868e-05 | V |  |  | 0.00069261 | V | 2 | 0.0013852 | V | 0.0001
10 | V | 5 | V | 50 | Hz | 5 | V | 0.00059717 | V | 0.00059717 | V | 0 | V | 2.8868e-06 | V | 0.00047627 | V | 0.00015011 | V | 2.3094e-05 | V | 2.8868e-05 | V |  |  | 0.00050074 | V | 2 | 0.0010015 | V | 0.0001
10 | V | 5 | V | 1 | kHz | 5 | V | 0.00059717 | V | 0.00059717 | V | 0 | V | 2.8868e-06 | V | 0.00047627 | V | 0.00015011 | V | 2.3094e-05 | V | 2.8868e-05 | V |  |  | 0.00050074 | V | 2 | 0.0010015 | V | 0.0001
10 | V | 5 | V | 10 | kHz | 5 | V | -0.00039531 | V | -0.00039531 | V | 0 | V | 2.8868e-06 | V | 0.0013948 | V | 0.00015011 | V | 2.3094e-05 | V | 2.8868e-05 | V |  |  | 0.0014034 | V | 2 | 0.0028068 | V | 0.0001
10 | V | 5 | V | 20 | kHz | 5 | V | -0.0009867 | V | -0.0009867 | V | 0 | V | 2.8868e-06 | V | 0.0040682 | V | 0.00025115 | V | 8.0829e-05 | V | 2.8868e-05 | V |  |  | 0.0040768 | V | 2 | 0.0081537 | V | 0.0001
10 | V | 5 | V | 50 | kHz | 5 | V | -0.0009867 | V | -0.0009867 | V | 0 | V | 2.8868e-06 | V | 0.0040682 | V | 0.00035507 | V | 0.00010392 | V | 2.8868e-05 | V |  |  | 0.0040851 | V | 2 | 0.0081702 | V | 0.0001
10 | V | 5 | V | 20 | kHz | 5 | V | -0.0009867 | V | -0.0009867 | V | 0 | V | 2.8868e-06 | V | 0.0040682 | V | 0.00025115 | V | 8.0829e-05 | V | 2.8868e-05 | V |  |  | 0.0040768 | V | 2 | 0.0081537 | V | 0.0001
10 | V | 10 | V | 20 | Hz | 10 | V | 0.00064346 | V | 0.00064346 | V | 0 | V | 2.8868e-06 | V | 0.00058666 | V | 0.00060622 | V | 0.00020785 | V | 2.8868e-05 | V |  |  | 0.00086931 | V | 2 | 0.0017386 | V | 0.0001
10 | V | 10 | V | 50 | Hz | 10 | V | 0.00095745 | V | 0.00095745 | V | 0 | V | 2.8868e-06 | V | 0.00047627 | V | 0.00027135 | V | 3.4641e-05 | V | 2.8868e-05 | V |  |  | 0.00055001 | V | 2 | 0.0011 | V | 0.0001
10 | V | 10 | V | 1 | kHz | 10 | V | 0.00095745 | V | 0.00095745 | V | 0 | V | 2.8868e-06 | V | 0.00047627 | V | 0.00027135 | V | 3.4641e-05 | V | 2.8868e-05 | V |  |  | 0.00055001 | V | 2 | 0.0011 | V | 0.0001
10 | V | 10 | V | 10 | kHz | 10 | V | -0.00087322 | V | -0.00087322 | V | 0 | V | 2.8868e-06 | V | 0.0013948 | V | 0.00027135 | V | 3.4641e-05 | V | 2.8868e-05 | V |  |  | 0.0014217 | V | 2 | 0.0028434 | V | 0.0001
10 | V | 10 | V | 20 | kHz | 10 | V | -0.0021324 | V | -0.0021324 | V | 0 | V | 2.8868e-06 | V | 0.0040682 | V | 0.00044456 | V | 0.00013856 | V | 2.8868e-05 | V |  |  | 0.0040949 | V | 2 | 0.0081897 | V | 0.0001
10 | V | 10 | V | 50 | kHz | 10 | V | -0.0021324 | V | -0.0021324 | V | 0 | V | 2.8868e-06 | V | 0.0040682 | V | 0.00059467 | V | 0.00016166 | V | 2.8868e-05 | V |  |  | 0.0041147 | V | 2 | 0.0082294 | V | 0.0001
10 | V | 10 | V | 100 | kHz | 10 | V | -0.003822 | V | -0.003822 | V | 0 | V | 2.8868e-06 | V | 0.0040691 | V | 0.0018129 | V | 0.001097 | V | 2.8868e-05 | V |  |  | 0.0045878 | V | 2 | 0.0091757 | V | 0.0001
100 | V | 10 | V | 20 | Hz | 10 | V | 0.00064346 | V | 0.00064346 | V | 0 | V | 2.8868e-05 | V | 0.00058666 | V | 0.00060622 | V | 0.00020785 | V | 2.8868e-05 | V |  |  | 0.00086979 | V | 2 | 0.0017396 | V | 0.0001
100 | V | 10 | V | 50 | Hz | 10 | V | 0.00095745 | V | 0.00095745 | V | 0 | V | 2.8868e-05 | V | 0.00047627 | V | 0.00027135 | V | 3.4641e-05 | V | 2.8868e-05 | V |  |  | 0.00055076 | V | 2 | 0.0011015 | V | 0.0001
100 | V | 10 | V | 1 | kHz | 10 | V | 0.00095745 | V | 0.00095745 | V | 0 | V | 2.8868e-05 | V | 0.00047627 | V | 0.00027135 | V | 3.4641e-05 | V | 2.8868e-05 | V |  |  | 0.00055076 | V | 2 | 0.0011015 | V | 0.0001
100 | V | 10 | V | 10 | kHz | 10 | V | -0.00087322 | V | -0.00087322 | V | 0 | V | 2.8868e-05 | V | 0.0013948 | V | 0.00027135 | V | 3.4641e-05 | V | 2.8868e-05 | V |  |  | 0.001422 | V | 2 | 0.002844 | V | 0.0001
100 | V | 10 | V | 20 | kHz | 10 | V | -0.0021324 | V | -0.0021324 | V | 0 | V | 2.8868e-05 | V | 0.0040682 | V | 0.00044456 | V | 0.00013856 | V | 2.8868e-05 | V |  |  | 0.004095 | V | 2 | 0.0081899 | V | 0.0001
100 | V | 10 | V | 50 | kHz | 10 | V | -0.0021324 | V | -0.0021324 | V | 0 | V | 2.8868e-05 | V | 0.0040682 | V | 0.00059467 | V | 0.00016166 | V | 2.8868e-05 | V |  |  | 0.0041148 | V | 2 | 0.0082296 | V | 0.0001
100 | V | 10 | V | 100 | kHz | 10 | V | -0.003822 | V | -0.003822 | V | 0 | V | 2.8868e-05 | V | 0.0040691 | V | 0.0018129 | V | 0.001097 | V | 2.8868e-05 | V |  |  | 0.0045879 | V | 2 | 0.0091758 | V | 0.0001
100 | V | 100 | V | 20 | Hz | 100 | V | -0.00028637 | V | -0.00028637 | V | 0 | V | 2.8868e-05 | V | 0.0058721 | V | 0.0060622 | V | 0.0020785 | V | 0.00028868 | V |  |  | 0.0086969 | V | 2 | 0.017394 | V | 0.001
100 | V | 100 | V | 50 | Hz | 100 | V | -0.0025353 | V | -0.0025353 | V | 0 | V | 2.8868e-05 | V | 0.0047512 | V | 0.0034641 | V | 0.00032332 | V | 0.00028868 | V |  |  | 0.005896 | V | 2 | 0.011792 | V | 0.001
100 | V | 100 | V | 1 | kHz | 100 | V | -0.0025353 | V | -0.0025353 | V | 0 | V | 2.8868e-05 | V | 0.0047512 | V | 0.0034641 | V | 0.00032332 | V | 0.00028868 | V |  |  | 0.005896 | V | 2 | 0.011792 | V | 0.001
100 | V | 100 | V | 10 | kHz | 100 | V | -0.029216 | V | -0.029216 | V | 0 | V | 2.8868e-05 | V | 0.01395 | V | 0.0034641 | V | 0.00032332 | V | 0.00028868 | V |  |  | 0.01438 | V | 2 | 0.028761 | V | 0.001
100 | V | 100 | V | 20 | kHz | 100 | V | -0.041907 | V | -0.041907 | V | 0 | V | 2.8868e-05 | V | 0.040973 | V | 0.0040415 | V | 0.0012702 | V | 0.00028868 | V |  |  | 0.041193 | V | 2 | 0.082385 | V | 0.001
100 | V | 100 | V | 50 | kHz | 100 | V | -0.041907 | V | -0.041907 | V | 0 | V | 2.8868e-05 | V | 0.040973 | V | 0.010104 | V | 0.0017321 | V | 0.00028868 | V |  |  | 0.042237 | V | 2 | 0.084474 | V | 0.001
100 | V | 100 | V | 100 | kHz | 100 | V | -0.040392 | V | -0.040392 | V | 0 | V | 2.8868e-05 | V | 0.041051 | V | 0.061199 | V | 0.0099304 | V | 0.0028868 | V |  |  | 0.074414 | V | 2 | 0.14883 | V | 0.01
1000 | V | 100 | V | 20 | Hz | 100 | V | -0.00028637 | V | -0.00028637 | V | 0 | V | 0.00028868 | V | 0.0058721 | V | 0.0060622 | V | 0.0020785 | V | 0.00028868 | V |  |  | 0.0087016 | V | 2 | 0.017403 | V | 0.001
1000 | V | 100 | V | 50 | Hz | 100 | V | -0.0025353 | V | -0.0025353 | V | 0 | V | 0.00028868 | V | 0.0047512 | V | 0.0034641 | V | 0.00032332 | V | 0.00028868 | V |  |  | 0.005903 | V | 2 | 0.011806 | V | 0.001
1000 | V | 100 | V | 1 | kHz | 100 | V | -0.0025353 | V | -0.0025353 | V | 0 | V | 0.00028868 | V | 0.0047512 | V | 0.0034641 | V | 0.00032332 | V | 0.00028868 | V |  |  | 0.005903 | V | 2 | 0.011806 | V | 0.001
1000 | V | 500 | V | 20 | Hz | 500 | V | 0.071289 | V | 0.071289 | V | 0 | V | 0 | V | 0.035411 | V | 0.09584 | V | 0.028868 | V | 0.00028868 | V |  |  | 0.10617 | V | 2 | 0.21235 | V | 0.001
1000 | V | 500 | V | 50 | Hz | 500 | V | 0.071289 | V | 0.071289 | V | 0 | V | 0.00028868 | V | 0.035411 | V | 0.013568 | V | 0.0011547 | V | 0.00028868 | V |  |  | 0.037941 | V | 2 | 0.075883 | V | 0.001
1000 | V | 500 | V | 1 | kHz | 500 | V | 0.071289 | V | 0.071289 | V | 0 | V | 0.00028868 | V | 0.035411 | V | 0.013568 | V | 0.0011547 | V | 0.00028868 | V |  |  | 0.037941 | V | 2 | 0.075883 | V | 0.001
1000 | V | 900 | V | 20 | Hz | 900 | V | 0.10968 | V | 0.10968 | V | 0 | V | 0 | V | 0.059727 | V | 0.16512 | V | 0.051962 | V | 0.00028868 | V |  |  | 0.18312 | V | 2 | 0.36624 | V | 0.001
1000 | V | 900 | V | 50 | Hz | 900 | V | 0.10968 | V | 0.10968 | V | 0 | V | 0.00028868 | V | 0.059727 | V | 0.022805 | V | 0.0020785 | V | 0.0028868 | V |  |  | 0.064032 | V | 2 | 0.12806 | V | 0.01
1000 | V | 900 | V | 1 | kHz | 900 | V | 0.10968 | V | 0.10968 | V | 0 | V | 0.00028868 | V | 0.059727 | V | 0.022805 | V | 0.0020785 | V | 0.0028868 | V |  |  | 0.064032 | V | 2 | 0.12806 | V | 0.01
Arus AC
Rentang UUC |  | Penunjukan Standar |  |  |  | Pembacaan alat |  | Koreksi Standar |  | Koreksi alat |  | ESDM UUC |  | RESOLUSI UUC |  | SERTIFIKAT STD |  | SPEK STD |  | TEMP STD |  | ROUNDING |  |  |  | Comb Unc |  | k | Exp Unc
100 | µA | 10 | µA | 20 | Hz | 10 | µA | -0.0024614 | µA | -0.0024614 | µA | 0 | µA | 2.8868e-05 | µA | 0.060077 | µA | 0.0066973 | µA | 0.0060044 | µA | 0.0028868 | µA |  |  | 0.060816 | µA | 2 | 0.12163 | µA | 0.01
100 | µA | 10 | µA | 50 | Hz | 10 | µA | 0.00042726 | µA | 0.00042726 | µA | 0 | µA | 2.8868e-05 | µA | 0.060077 | µA | 0.0052135 | µA | 0.00062354 | µA | 0.0028868 | µA |  |  | 0.060375 | µA | 2 | 0.12075 | µA | 0.01
100 | µA | 10 | µA | 1 | kHz | 10 | µA | 0.00042726 | µA | 0.00042726 | µA | 0 | µA | 2.8868e-05 | µA | 0.060077 | µA | 0.0085448 | µA | 0.0012702 | µA | 0.0028868 | µA |  |  | 0.060763 | µA | 2 | 0.12153 | µA | 0.01
100 | µA | 100 | µA | 20 | Hz | 100 | µA | 0.0033687 | µA | 0.0033687 | µA | 0 | µA | 2.8868e-05 | µA | 0.061759 | µA | 0.015011 | µA | 0.0080829 | µA | 0.0028868 | µA |  |  | 0.064134 | µA | 2 | 0.12827 | µA | 0.01
100 | µA | 100 | µA | 50 | Hz | 100 | µA | 0.0033687 | µA | 0.0033687 | µA | 0 | µA | 2.8868e-05 | µA | 0.061754 | µA | 0.010566 | µA | 0.0010392 | µA | 0.0028868 | µA |  |  | 0.062727 | µA | 2 | 0.12545 | µA | 0.01
100 | µA | 100 | µA | 1 | kHz | 100 | µA | 0.0033687 | µA | 0.0033687 | µA | 0 | µA | 2.8868e-05 | µA | 0.061754 | µA | 0.023094 | µA | 0.0023094 | µA | 0.0028868 | µA |  |  | 0.066035 | µA | 2 | 0.13207 | µA | 0.01
1 | mA | 0.1 | mA | 20 | Hz | 0.1 | mA | 2.2131e-06 | mA | 2.2131e-06 | mA | 0 | mA | 2.8868e-07 | mA | 6.1759e-05 | mA | 1.5011e-05 | mA | 8.0829e-06 | mA | 2.8868e-06 | mA |  |  | 6.4135e-05 | mA | 2 | 0.00012827 | mA | 1e-05
1 | mA | 0.1 | mA | 50 | Hz | 0.1 | mA | 3.3687e-06 | mA | 3.3687e-06 | mA | 0 | mA | 2.8868e-07 | mA | 6.1754e-05 | mA | 1.0566e-05 | mA | 1.0392e-06 | mA | 2.8868e-06 | mA |  |  | 6.2727e-05 | mA | 2 | 0.00012545 | mA | 1e-05
1 | mA | 0.1 | mA | 1 | kHz | 0.1 | mA | 3.3687e-06 | mA | 3.3687e-06 | mA | 0 | mA | 2.8868e-07 | mA | 6.1754e-05 | mA | 2.3094e-05 | mA | 2.3094e-06 | mA | 2.8868e-06 | mA |  |  | 6.6035e-05 | mA | 2 | 0.00013207 | mA | 1e-05
1 | mA | 1 | mA | 20 | Hz | 1 | mA | 9.3378e-05 | mA | 9.3378e-05 | mA | 0 | mA | 2.8868e-07 | mA | 0.00026722 | mA | 0.00011258 | mA | 2.7713e-05 | mA | 2.8868e-06 | mA |  |  | 0.00029131 | mA | 2 | 0.00058261 | mA | 1e-05
1 | mA | 1 | mA | 50 | Hz | 1 | mA | 8.6831e-05 | mA | 8.6831e-05 | mA | 0 | mA | 2.8868e-07 | mA | 0.00026707 | mA | 7.9674e-05 | mA | 5.7735e-06 | mA | 2.8868e-06 | mA |  |  | 0.00027878 | mA | 2 | 0.00055756 | mA | 1e-05
1 | mA | 1 | mA | 1 | kHz | 1 | mA | 8.6831e-05 | mA | 8.6831e-05 | mA | 0 | mA | 2.8868e-07 | mA | 0.00026707 | mA | 0.00017898 | mA | 0.00012702 | mA | 2.8868e-06 | mA |  |  | 0.00034569 | mA | 2 | 0.00069139 | mA | 1e-05
1 | mA | 1 | mA | 5 | kHz | 1 | mA | 6.5019e-05 | mA | 6.5019e-05 | mA | 0 | mA | 2.8868e-07 | mA | 0.00026708 | mA | 0.0010104 | mA | 0.00051962 | mA | 2.8868e-05 | mA |  |  | 0.0011675 | mA | 2 | 0.0023349 | mA | 0.0001
1 | mA | 1 | mA | 10 | kHz | 1 | mA | -0.00016324 | mA | -0.00016324 | mA | 0 | mA | 2.8868e-07 | mA | 0.00027875 | mA | 0.0010104 | mA | 0.00051962 | mA | 2.8868e-05 | mA |  |  | 0.0011702 | mA | 2 | 0.0023404 | mA | 0.0001
10 | mA | 1 | mA | 20 | Hz | 1 | mA | 9.3378e-05 | mA | 9.3378e-05 | mA | 0 | mA | 2.8868e-06 | mA | 0.00026722 | mA | 0.00011258 | mA | 2.7713e-05 | mA | 2.8868e-06 | mA |  |  | 0.00029132 | mA | 2 | 0.00058264 | mA | 1e-05
10 | mA | 1 | mA | 50 | Hz | 1 | mA | 8.6831e-05 | mA | 8.6831e-05 | mA | 0 | mA | 2.8868e-06 | mA | 0.00026707 | mA | 7.9674e-05 | mA | 5.7735e-06 | mA | 2.8868e-06 | mA |  |  | 0.0002788 | mA | 2 | 0.00055759 | mA | 1e-05
10 | mA | 1 | mA | 1 | kHz | 1 | mA | 8.6831e-05 | mA | 8.6831e-05 | mA | 0 | mA | 2.8868e-06 | mA | 0.00026707 | mA | 0.00017898 | mA | 0.00012702 | mA | 2.8868e-06 | mA |  |  | 0.00034571 | mA | 2 | 0.00069141 | mA | 1e-05
10 | mA | 1 | mA | 5 | kHz | 1 | mA | 6.5019e-05 | mA | 6.5019e-05 | mA | 0 | mA | 2.8868e-06 | mA | 0.00026708 | mA | 0.0010104 | mA | 0.00051962 | mA | 2.8868e-05 | mA |  |  | 0.0011675 | mA | 2 | 0.002335 | mA | 0.0001
10 | mA | 1 | mA | 10 | kHz | 1 | mA | -0.00016324 | mA | -0.00016324 | mA | 0 | mA | 2.8868e-06 | mA | 0.00027875 | mA | 0.0010104 | mA | 0.00051962 | mA | 2.8868e-05 | mA |  |  | 0.0011702 | mA | 2 | 0.0023404 | mA | 0.0001
10 | mA | 5 | mA | 20 | Hz | 5 | mA | 0.00071403 | mA | 0.00071403 | mA | 0 | mA | 2.8868e-06 | mA | 0.0019719 | mA | 0.00066395 | mA | 0.00012702 | mA | 2.8868e-05 | mA |  |  | 0.0020848 | mA | 2 | 0.0041696 | mA | 0.0001
10 | mA | 5 | mA | 50 | Hz | 5 | mA | 0.00029322 | mA | 0.00029322 | mA | 0 | mA | 2.8868e-06 | mA | 0.0019717 | mA | 0.00049941 | mA | 3.4641e-05 | mA | 2.8868e-05 | mA |  |  | 0.0020345 | mA | 2 | 0.0040689 | mA | 0.0001
10 | mA | 5 | mA | 1 | kHz | 5 | mA | 0.00029322 | mA | 0.00029322 | mA | 0 | mA | 2.8868e-06 | mA | 0.0019717 | mA | 0.00089489 | mA | 0.00063509 | mA | 2.8868e-05 | mA |  |  | 0.0022567 | mA | 2 | 0.0045134 | mA | 0.0001
10 | mA | 5 | mA | 5 | kHz | 5 | mA | 0.00052722 | mA | 0.00052722 | mA | 0 | mA | 2.8868e-06 | mA | 0.0019717 | mA | 0.0060622 | mA | 0.0014434 | mA | 0.00028868 | mA |  |  | 0.0065425 | mA | 2 | 0.013085 | mA | 0.001
10 | mA | 5 | mA | 10 | kHz | 5 | mA | 0.00046212 | mA | 0.00046212 | mA | 0 | mA | 2.8868e-06 | mA | 0.0021278 | mA | 0.0060622 | mA | 0.0014434 | mA | 0.00028868 | mA |  |  | 0.0065912 | mA | 2 | 0.013182 | mA | 0.001
10 | mA | 10 | mA | 20 | Hz | 10 | mA | 0.0010566 | mA | 0.0010566 | mA | 0 | mA | 2.8868e-06 | mA | 0.0026715 | mA | 0.0011258 | mA | 0.00024249 | mA | 2.8868e-05 | mA |  |  | 0.0029093 | mA | 2 | 0.0058186 | mA | 0.0001
10 | mA | 10 | mA | 50 | Hz | 10 | mA | 0.00073198 | mA | 0.00073198 | mA | 0 | mA | 2.8868e-06 | mA | 0.0026708 | mA | 0.00079674 | mA | 5.7735e-05 | mA | 2.8868e-05 | mA |  |  | 0.0027878 | mA | 2 | 0.0055757 | mA | 0.0001
10 | mA | 10 | mA | 1 | kHz | 10 | mA | 0.00073198 | mA | 0.00073198 | mA | 0 | mA | 2.8868e-06 | mA | 0.0026708 | mA | 0.0014722 | mA | 0.00069282 | mA | 2.8868e-05 | mA |  |  | 0.0031275 | mA | 2 | 0.006255 | mA | 0.0001
10 | mA | 10 | mA | 5 | kHz | 10 | mA | 0.00078076 | mA | 0.00078076 | mA | 0 | mA | 2.8868e-06 | mA | 0.0026708 | mA | 0.0092376 | mA | 0.0017321 | mA | 0.00028868 | mA |  |  | 0.009775 | mA | 2 | 0.01955 | mA | 0.001
10 | mA | 10 | mA | 10 | kHz | 10 | mA | 0.00015365 | mA | 0.00015365 | mA | 0 | mA | 2.8868e-06 | mA | 0.0027878 | mA | 0.0092376 | mA | 0.0017321 | mA | 0.00028868 | mA |  |  | 0.0098076 | mA | 2 | 0.019615 | mA | 0.001
100 | mA | 10 | mA | 20 | Hz | 10 | mA | 0.0010566 | mA | 0.0010566 | mA | 0 | mA | 2.8868e-05 | mA | 0.0026715 | mA | 0.0011258 | mA | 0.00024249 | mA | 2.8868e-05 | mA |  |  | 0.0029095 | mA | 2 | 0.0058189 | mA | 0.0001
100 | mA | 10 | mA | 50 | Hz | 10 | mA | 0.00073198 | mA | 0.00073198 | mA | 0 | mA | 2.8868e-05 | mA | 0.0026708 | mA | 0.00079674 | mA | 5.7735e-05 | mA | 2.8868e-05 | mA |  |  | 0.002788 | mA | 2 | 0.005576 | mA | 0.0001
100 | mA | 10 | mA | 1 | kHz | 10 | mA | 0.00073198 | mA | 0.00073198 | mA | 0 | mA | 2.8868e-05 | mA | 0.0026708 | mA | 0.0014722 | mA | 0.00069282 | mA | 2.8868e-05 | mA |  |  | 0.0031277 | mA | 2 | 0.0062553 | mA | 0.0001
100 | mA | 10 | mA | 5 | kHz | 10 | mA | 0.00078076 | mA | 0.00078076 | mA | 0 | mA | 2.8868e-05 | mA | 0.0026708 | mA | 0.0092376 | mA | 0.0017321 | mA | 0.00028868 | mA |  |  | 0.009775 | mA | 2 | 0.01955 | mA | 0.001
100 | mA | 10 | mA | 10 | kHz | 10 | mA | 0.00015365 | mA | 0.00015365 | mA | 0 | mA | 2.8868e-05 | mA | 0.0027878 | mA | 0.0092376 | mA | 0.0017321 | mA | 0.00028868 | mA |  |  | 0.0098076 | mA | 2 | 0.019615 | mA | 0.001
100 | mA | 100 | mA | 20 | Hz | 100 | mA | 0.0070198 | mA | 0.0070198 | mA | 0 | mA | 2.8868e-05 | mA | 0.026748 | mA | 0.011258 | mA | 0.0023671 | mA | 0.00028868 | mA |  |  | 0.029118 | mA | 2 | 0.058236 | mA | 0.001
100 | mA | 100 | mA | 50 | Hz | 100 | mA | 0.0074181 | mA | 0.0074181 | mA | 0 | mA | 2.8868e-05 | mA | 0.02674 | mA | 0.0073901 | mA | 0.00057735 | mA | 0.00028868 | mA |  |  | 0.02775 | mA | 2 | 0.0555 | mA | 0.001
100 | mA | 100 | mA | 1 | kHz | 100 | mA | 0.0074181 | mA | 0.0074181 | mA | 0 | mA | 2.8868e-05 | mA | 0.02674 | mA | 0.013568 | mA | 0.0034641 | mA | 0.00028868 | mA |  |  | 0.030186 | mA | 2 | 0.060372 | mA | 0.001
100 | mA | 100 | mA | 5 | kHz | 100 | mA | 0.012396 | mA | 0.012396 | mA | 0 | mA | 2.8868e-05 | mA | 0.02674 | mA | 0.069282 | mA | 0.011547 | mA | 0.0028868 | mA |  |  | 0.075211 | mA | 2 | 0.15042 | mA | 0.01
100 | mA | 100 | mA | 10 | kHz | 100 | mA | -0.002865 | mA | -0.002865 | mA | 0 | mA | 2.8868e-05 | mA | 0.027909 | mA | 0.069282 | mA | 0.011547 | mA | 0.0028868 | mA |  |  | 0.075635 | mA | 2 | 0.15127 | mA | 0.01
1 | A | 0.1 | A | 20 | Hz | 0.1 | A | 7.0198e-06 | A | 7.0198e-06 | A | 0 | A | 2.8868e-07 | A | 2.6748e-05 | A | 1.1258e-05 | A | 2.3671e-06 | A | 2.8868e-07 | A |  |  | 2.912e-05 | A | 2 | 5.8239e-05 | A | 1e-06
1 | A | 0.1 | A | 50 | Hz | 0.1 | A | 7.4181e-06 | A | 7.4181e-06 | A | 0 | A | 2.8868e-07 | A | 2.674e-05 | A | 7.3901e-06 | A | 5.7735e-07 | A | 2.8868e-07 | A |  |  | 2.7751e-05 | A | 2 | 5.5503e-05 | A | 1e-06
1 | A | 0.1 | A | 1 | kHz | 0.1 | A | 7.4181e-06 | A | 7.4181e-06 | A | 0 | A | 2.8868e-07 | A | 2.674e-05 | A | 1.3568e-05 | A | 3.4641e-06 | A | 2.8868e-07 | A |  |  | 3.0187e-05 | A | 2 | 6.0375e-05 | A | 1e-06
1 | A | 0.1 | A | 5 | kHz | 0.1 | A | 1.2396e-05 | A | 1.2396e-05 | A | 0 | A | 2.8868e-07 | A | 2.674e-05 | A | 6.9282e-05 | A | 1.1547e-05 | A | 2.8868e-06 | A |  |  | 7.5212e-05 | A | 2 | 0.00015042 | A | 1e-05
1 | A | 0.1 | A | 10 | kHz | 0.1 | A | -2.865e-06 | A | -2.865e-06 | A | 0 | A | 2.8868e-07 | A | 2.7909e-05 | A | 6.9282e-05 | A | 1.1547e-05 | A | 2.8868e-06 | A |  |  | 7.5635e-05 | A | 2 | 0.00015127 | A | 1e-05
1 | A | 1 | A | 20 | Hz | 1 | A | -1.162e-05 | A | -1.162e-05 | A | 0 | A | 2.8868e-07 | A | 0.00047774 | A | 0.00016108 | A | 5.7735e-06 | A | 2.8868e-05 | A |  |  | 0.00050502 | A | 2 | 0.00101 | A | 0.0001
1 | A | 1 | A | 50 | Hz | 1 | A | -2.5673e-06 | A | -2.5673e-06 | A | 0 | A | 2.8868e-07 | A | 0.00047779 | A | 0.00016108 | A | 5.7735e-06 | A | 2.8868e-05 | A |  |  | 0.00050507 | A | 2 | 0.0010101 | A | 0.0001
1 | A | 1 | A | 1 | kHz | 1 | A | -2.5673e-06 | A | -2.5673e-06 | A | 0 | A | 2.8868e-07 | A | 0.00047779 | A | 0.000306 | A | 1.7321e-05 | A | 2.8868e-05 | A |  |  | 0.00056837 | A | 2 | 0.0011367 | A | 0.0001
1 | A | 1 | A | 5 | kHz | 1 | A | 0.00027853 | A | 0.00027853 | A | 0 | A | 2.8868e-07 | A | 0.00053421 | A | 0.0041338 | A | 6.9282e-05 | A | 2.8868e-05 | A |  |  | 0.0041689 | A | 2 | 0.0083378 | A | 0.0001
1 | A | 1 | A | 10 | kHz | 1 | A | 0.0011969 | A | 0.0011969 | A | 0 | A | 2.8868e-07 | A | 0.00055786 | A | 0.0041338 | A | 6.9282e-05 | A | 2.8868e-05 | A |  |  | 0.004172 | A | 2 | 0.008344 | A | 0.0001
Resistansi
Rentang UUC |  | Penunjukan Standar |  |  |  | Pembacaan alat |  | Koreksi Standar |  | Koreksi alat |  | ESDM UUC |  | RESOLUSI UUC |  | SERTIFIKAT STD |  | SPEK STD |  | TEMP STD |  | ROUNDING |  |  |  | Comb Unc |  | k | Exp Unc
10 | Ω |  | 1 | Ω |  | 1 | Ω | 2.27e-05 | Ω | 2.27e-05 | Ω | 0 | Ω | 2.8868e-06 | Ω | 1.15e-06 | Ω | 4.6188e-06 | Ω | 3.4641e-06 | Ω | 2.8868e-06 | Ω |  |  | 7.164e-06 | Ω | 2 | 1.4328e-05 | Ω | 1e-05
10 | Ω |  | 10 | Ω |  | 10 | Ω | 8e-06 | Ω | 8e-06 | Ω | 0 | Ω | 2.8868e-06 | Ω | 1.15e-05 | Ω | 4.6188e-05 | Ω | 3.4641e-05 | Ω | 2.8868e-06 | Ω |  |  | 5.9011e-05 | Ω | 2 | 0.00011802 | Ω | 1e-05
100 | Ω |  | 10 | Ω |  | 10 | Ω | 8e-06 | Ω | 8e-06 | Ω | 0 | Ω | 2.8868e-06 | Ω | 1.15e-05 | Ω | 4.6188e-05 | Ω | 3.4641e-05 | Ω | 2.8868e-06 | Ω |  |  | 5.9011e-05 | Ω | 2 | 0.00011802 | Ω | 1e-05
100 | Ω |  | 100 | Ω |  | 100 | Ω | 0.0005 | Ω | 0.0005 | Ω | 0 | Ω | 2.8868e-06 | Ω | 0.000115 | Ω | 0.00034641 | Ω | 0.00034641 | Ω | 2.8868e-05 | Ω |  |  | 0.00050405 | Ω | 2 | 0.0010081 | Ω | 0.0001
1 | kΩ |  | 0.1 | kΩ |  | 0.1 | kΩ | 5e-07 | kΩ | 5e-07 | kΩ | 0 | kΩ | 2.8868e-08 | kΩ | 1.15e-07 | kΩ | 3.4641e-07 | kΩ | 3.4641e-07 | kΩ | 2.8868e-08 | kΩ |  |  | 5.0487e-07 | kΩ | 2 | 1.0097e-06 | kΩ | 1e-07
1 | kΩ |  | 1 | kΩ |  | 1 | kΩ | 5.71e-05 | kΩ | 5.71e-05 | kΩ | 0 | kΩ | 2.8868e-08 | kΩ | 1.15e-06 | kΩ | 3.4641e-06 | kΩ | 2.3094e-06 | kΩ | 2.8868e-05 | kΩ |  |  | 2.9189e-05 | kΩ | 2 | 5.8378e-05 | kΩ | 0.0001
10 | kΩ |  | 1 | kΩ |  | 1 | kΩ | 5.71e-05 | kΩ | 5.71e-05 | kΩ | 0 | kΩ | 2.8868e-07 | kΩ | 1.15e-06 | kΩ | 3.4641e-06 | kΩ | 2.3094e-06 | kΩ | 2.8868e-05 | kΩ |  |  | 2.919e-05 | kΩ | 2 | 5.8381e-05 | kΩ | 0.0001
10 | kΩ |  | 10 | kΩ |  | 10 | kΩ | 5.4e-05 | kΩ | 5.4e-05 | kΩ | 0 | kΩ | 2.8868e-07 | kΩ | 1.15e-05 | kΩ | 2.3094e-05 | kΩ | 1.7321e-05 | kΩ | 2.8868e-07 | kΩ |  |  | 3.1077e-05 | kΩ | 2 | 6.2153e-05 | kΩ | 1e-06
100 | kΩ |  | 10 | kΩ |  | 10 | kΩ | 5.4e-05 | kΩ | 5.4e-05 | kΩ | 0 | kΩ | 2.8868e-06 | kΩ | 1.15e-05 | kΩ | 2.3094e-05 | kΩ | 1.7321e-05 | kΩ | 2.8868e-06 | kΩ |  |  | 3.1341e-05 | kΩ | 2 | 6.2682e-05 | kΩ | 1e-05
100 | kΩ |  | 100 | kΩ |  | 100 | kΩ | 0.00085 | kΩ | 0.00085 | kΩ | 0 | kΩ | 2.8868e-06 | kΩ | 0.000365 | kΩ | 0.00034641 | kΩ | 0.00023094 | kΩ | 2.8868e-05 | kΩ |  |  | 0.00055444 | kΩ | 2 | 0.0011089 | kΩ | 0.0001
1 | MΩ |  | 0.1 | MΩ |  | 0.1 | MΩ | 8.5e-07 | MΩ | 8.5e-07 | MΩ | 0 | MΩ | 2.8868e-08 | MΩ | 3.65e-07 | MΩ | 3.4641e-07 | MΩ | 2.3094e-07 | MΩ | 2.8868e-08 | MΩ |  |  | 5.5518e-07 | MΩ | 2 | 1.1104e-06 | MΩ | 1e-07
1 | MΩ |  | 1 | MΩ |  | 1 | MΩ | 4e-06 | MΩ | 4e-06 | MΩ | 0 | MΩ | 2.8868e-08 | MΩ | 5.5e-06 | MΩ | 4.6188e-06 | MΩ | 2.3094e-06 | MΩ | 2.8868e-07 | MΩ |  |  | 7.5499e-06 | MΩ | 2 | 1.51e-05 | MΩ | 1e-06
10 | MΩ |  | 1 | MΩ |  | 1 | MΩ | 4e-06 | MΩ | 4e-06 | MΩ | 0 | MΩ | 2.8868e-07 | MΩ | 5.5e-06 | MΩ | 4.6188e-06 | MΩ | 2.3094e-06 | MΩ | 2.8868e-07 | MΩ |  |  | 7.5554e-06 | MΩ | 2 | 1.5111e-05 | MΩ | 1e-06
10 | MΩ |  | 10 | MΩ |  | 10 | MΩ | 8e-05 | MΩ | 8e-05 | MΩ | 0 | MΩ | 2.8868e-07 | MΩ | 0.000105 | MΩ | 5.7735e-05 | MΩ | 5.7735e-05 | MΩ | 2.8868e-06 | MΩ |  |  | 0.00013304 | MΩ | 2 | 0.00026608 | MΩ | 1e-05
100 | MΩ |  | 10 | MΩ |  | 10 | MΩ | 8e-05 | MΩ | 8e-05 | MΩ | 0 | MΩ | 2.8868e-06 | MΩ | 0.000105 | MΩ | 5.7735e-05 | MΩ | 5.7735e-05 | MΩ | 2.8868e-06 | MΩ |  |  | 0.00013307 | MΩ | 2 | 0.00026615 | MΩ | 1e-05
100 | MΩ |  | 100 | MΩ |  | 100 | MΩ | 0.0131 | MΩ | 0.0131 | MΩ | 0 | MΩ | 2.8868e-06 | MΩ | 0.00115 | MΩ | 0.0014434 | MΩ | 0.00057735 | MΩ | 2.8868e-05 | MΩ |  |  | 0.0019339 | MΩ | 2 | 0.0038678 | MΩ | 0.0001
1 | GΩ |  | 0.1 | GΩ |  | 0.1 | GΩ | 1.31e-05 | GΩ | 1.31e-05 | GΩ | 0 | GΩ | 2.8868e-08 | GΩ | 1.15e-06 | GΩ | 1.4434e-06 | GΩ | 5.7735e-07 | GΩ | 2.8868e-08 | GΩ |  |  | 1.9341e-06 | GΩ | 2 | 3.8682e-06 | GΩ | 1e-07
1 | GΩ |  | 1 | GΩ |  | 1 | GΩ | 0.001067 | GΩ | 0.001067 | GΩ | 0 | GΩ | 2.8868e-08 | GΩ | 1.9e-05 | GΩ | 2.0207e-05 | GΩ | 6.9282e-06 | GΩ | 2.8868e-06 | GΩ |  |  | 2.8734e-05 | GΩ | 2 | 5.7469e-05 | GΩ | 1e-05

## Lap
Nama Alat/Instrument Name | [7 empty] | : 8,5 Digit Multimeter
Pembuat/Manufacturer | [7 empty] | : Keysight
Model/Model | [7 empty] | : 3458A
No.Seri/Serial Number | [7 empty] | : MY12345678
Tanggal Kalibrasi/Calibration Date | [7 empty] | : 21 Februari - 5 Maret 2024
Tempat Kalibrasi/Calibration Place | [7 empty] | : Lab SNSU BSN
Hasil Kalibrasi/Calibration Result
Kondisi Ruangan/Environmental Condition
Suhu/Temperature | [7 empty] | : (22 ± 1) °C
Kelembaban Relatif/Relative Humidity | [7 empty] | : (50 ± 5) %RH
Tegangan DC / DC Voltage
Rentang |  |  | Titik Ukur |  |  |  | Pembacaan Alat |  | Koreksi |  | Ketidakpastian
Range |  |  | Measurement Point |  |  |  | Instrument Reading |  | Correction |  | Uncertainty
| 100 | mV |  | 0 | mV* |  | 0 | mV | -2.1073e-05 | mV | 0.001327 | mV
|  |  |  | 10 | mV |  | 10 | mV | -0.00011587 | mV | 0.0013643 | mV
|  |  |  | -10 | mV |  | -10 | mV | -2.3787e-05 | mV | 0.0013616 | mV
|  |  |  | 50 | mV |  | 50 | mV | -0.00049506 | mV | 0.0015742 | mV
|  |  |  | -50 | mV |  | -50 | mV | -7.1414e-05 | mV | 0.0015606 | mV
|  |  |  | 90 | mV |  | 90 | mV | -0.00087426 | mV | 0.0018244 | mV
|  |  |  | -90 | mV |  | -90 | mV | -0.00011904 | mV | 0.0018033 | mV
|  |  |  | 100 | mV |  | 100 | mV | -0.00096905 | mV | 0.001893 | mV
|  |  |  | -100 | mV |  | -100 | mV | -0.00013095 | mV | 0.0018703 | mV
| 1 | V |  | 0.1 | V |  | 0.1 | V | -9.6905e-07 | V | 1.8938e-06 | V
|  |  |  | -0.1 | V |  | -0.1 | V | -1.3095e-07 | V | 1.8712e-06 | V
|  |  |  | 0.5 | V |  | 0.5 | V | -1.6602e-06 | V | 4.683e-06 | V
|  |  |  | -0.5 | V |  | -0.5 | V | 6.4034e-07 | V | 4.65e-06 | V
|  |  |  | 0.9 | V |  | 0.9 | V | -2.0254e-06 | V | 7.182e-06 | V
|  |  |  | -0.9 | V |  | -0.9 | V | 5.2655e-07 | V | 7.1427e-06 | V
|  |  |  | 1 | V |  | 1 | V | -2.1167e-06 | V | 7.8228e-06 | V
|  |  |  | -1 | V |  | -1 | V | 4.981e-07 | V | 7.782e-06 | V
| 10 | V |  | 1 | V |  | 1 | V | -2.1167e-06 | V | 7.8648e-06 | V
|  |  |  | -1 | V |  | -1 | V | 4.981e-07 | V | 7.8243e-06 | V
|  |  |  | 2 | V |  | 2 | V | -3.0297e-06 | V | 1.4424e-05 | V
|  |  |  | 3 | V |  | 3 | V | -1.0427e-05 | V | 2.6751e-05 | V
|  |  |  | 4 | V |  | 4 | V | -1.1942e-05 | V | 2.923e-05 | V
|  |  |  | 5 | V |  | 5 | V | -1.3458e-05 | V | 3.2034e-05 | V
|  |  |  | -5 | V |  | -5 | V | 9.8766e-06 | V | 3.209e-05 | V
Tegangan DC / DC Voltage
Rentang |  |  | Titik Ukur |  |  |  | Pembacaan Alat |  | Koreksi |  | Ketidakpastian
Range |  |  | Measurement Point |  |  |  | Instrument Reading |  | Correction |  | Uncertainty
| 10 | V |  | 6 | V |  | 6 | V | -1.4973e-05 | V | 4.4978e-05 | V
|  |  |  | 7 | V |  | 7 | V | -1.6488e-05 | V | 4.7545e-05 | V
|  |  |  | 8 | V |  | 8 | V | -1.8003e-05 | V | 5.0309e-05 | V
|  |  |  | 9 | V |  | 9 | V | -1.9518e-05 | V | 5.3239e-05 | V
|  |  |  | -9 | V |  | -9 | V | 8.0177e-06 | V | 5.319e-05 | V
|  |  |  | 10 | V |  | 10 | V | -2.1033e-05 | V | 5.8236e-05 | V
|  |  |  | -10 | V |  | -10 | V | 7.553e-06 | V | 5.8188e-05 | V
| 100 | V |  | 10 | V |  | 10 | V | -2.1033e-05 | V | 5.88e-05 | V
|  |  |  | -10 | V |  | -10 | V | 7.553e-06 | V | 5.8753e-05 | V
|  |  |  | 50 | V |  | 50 | V | -3.2271e-05 | V | 0.0004607 | V
|  |  |  | -50 | V |  | -50 | V | 4.7284e-05 | V | 0.00045894 | V
|  |  |  | 90 | V |  | 90 | V | -5.2675e-05 | V | 0.00075809 | V
|  |  |  | -90 | V |  | -90 | V | 5.1501e-05 | V | 0.00075627 | V
|  |  |  | 100 | V |  | 100 | V | -5.7776e-05 | V | 0.00083319 | V
|  |  |  | -100 | V |  | -100 | V | 5.2556e-05 | V | 0.00083127 | V
| 1000 | V |  | 100 | V |  | 100 | V | -5.7776e-05 | V | 0.00083714 | V
|  |  |  | -100 | V |  | -100 | V | 5.2556e-05 | V | 0.00083523 | V
|  |  |  | 500 | V |  | 500 | V | -0.00063642 | V | 0.0042832 | V
|  |  |  | -500 | V |  | -500 | V | 0.00013062 | V | 0.0042819 | V
|  |  |  | 900 | V |  | 900 | V | -0.00076853 | V | 0.0077363 | V
|  |  |  | -900 | V |  | -900 | V | 0.00033487 | V | 0.0077341 | V
|  |  |  | 1000 | V |  | 1000 | V | -0.00080156 | V | 0.0081545 | V
|  |  |  | -1000 | V |  | -1000 | V | 0.00038594 | V | 0.0081518 | V
Arus DC / DC Current
Rentang |  |  | Titik Ukur |  |  |  | Pembacaan Alat |  | Koreksi |  | Ketidakpastian
Range |  |  | Measurement Point |  |  |  | Instrument Reading |  | Correction |  | Uncertainty
| 100 | µA |  | 0 | µA* |  | 0.001514 | µA | 0.0025852 | µA | 0.0070467 | µA
|  |  |  | 10 | µA |  | 9.9984 | µA | 0.005581 | µA | 0.0075164 | µA
|  |  |  | -10 | µA |  | -10.0013 | µA | 0.0055712 | µA | 0.0074917 | µA
|  |  |  | 50 | µA |  | 49.9976 | µA | 0.0058202 | µA | 0.0094149 | µA
|  |  |  | -50 | µA |  | -50.0005 | µA | 0.0052182 | µA | 0.009319 | µA
|  |  |  | 90 | µA |  | 89.9968 | µA | 0.0060974 | µA | 0.011333 | µA
|  |  |  | -90 | µA |  | -89.9998 | µA | 0.0049793 | µA | 0.011187 | µA
|  |  |  | 100 | µA |  | 99.9965 | µA | 0.0062492 | µA | 0.011815 | µA
|  |  |  | -100 | µA |  | -99.9996 | µA | 0.0049611 | µA | 0.011652 | µA
| 1 | mA |  | 0.1 | mA |  | 0.099992 | mA | 1.0597e-05 | mA | 1.2336e-05 | mA
|  |  |  | -0.1 | mA |  | -0.10002 | mA | 2.3453e-05 | mA | 1.4799e-05 | mA
|  |  |  | 0.5 | mA |  | 0.50002 | mA | -1.4319e-05 | mA | 3.1877e-05 | mA
Arus DC / DC Current
Rentang |  |  | Titik Ukur |  |  |  | Pembacaan Alat |  | Koreksi |  | Ketidakpastian
Range |  |  | Measurement Point |  |  |  | Instrument Reading |  | Correction |  | Uncertainty
| 1 | mA |  | -0.5 | mA |  | -0.50004 | mA | 4.8191e-05 | mA | 3.6554e-05 | mA
|  |  |  | 0.9 | mA |  | 0.90004 | mA | -3.8929e-05 | mA | 5.1629e-05 | mA
|  |  |  | -0.9 | mA |  | -0.90007 | mA | 7.1417e-05 | mA | 5.6548e-05 | mA
|  |  |  | 1 | mA |  | 1.0001 | mA | -4.4761e-05 | mA | 5.66e-05 | mA
|  |  |  | -1 | mA |  | -1.0001 | mA | 7.7983e-05 | mA | 6.179e-05 | mA
| 10 | mA |  | 1 | mA |  | 0.99988 | mA | 0.00012542 | mA | 7.8985e-05 | mA
|  |  |  | -1 | mA |  | -1.0001 | mA | 0.000153 | mA | 8.9155e-05 | mA
|  |  |  | 2 | mA |  | 1.9999 | mA | 0.0001119 | mA | 0.00010623 | mA
|  |  |  | 3 | mA |  | 2.9999 | mA | 8.8886e-05 | mA | 0.00020156 | mA
|  |  |  | 4 | mA |  | 4 | mA | 7.3743e-05 | mA | 0.00023585 | mA
|  |  |  | 5 | mA |  | 5 | mA | 5.46e-05 | mA | 0.0002724 | mA
|  |  |  | -5 | mA |  | -5.0002 | mA | 0.00020258 | mA | 0.0002881 | mA
|  |  |  | 6 | mA |  | 6 | mA | 4.4258e-05 | mA | 0.00032457 | mA
|  |  |  | 7 | mA |  | 7 | mA | 3.5715e-05 | mA | 0.00036174 | mA
|  |  |  | 8 | mA |  | 8 | mA | 2.0772e-05 | mA | 0.00039958 | mA
|  |  |  | 9 | mA |  | 9.0001 | mA | 8.629e-06 | mA | 0.00043818 | mA
|  |  |  | -9 | mA |  | -9.0003 | mA | 0.00025028 | mA | 0.00045603 | mA
|  |  |  | 10 | mA |  | 10.0001 | mA | -1.0914e-05 | mA | 0.00048029 | mA
|  |  |  | -10 | mA |  | -10.0003 | mA | 0.00027041 | mA | 0.00049959 | mA
| 100 | mA |  | 10 | mA |  | 10.0002 | mA | -7.6914e-05 | mA | 0.00049302 | mA
|  |  |  | -10 | mA |  | -10.0027 | mA | 0.0026632 | mA | 0.0014386 | mA
|  |  |  | 50 | mA |  | 50.0063 | mA | -0.0052086 | mA | 0.0050501 | mA
|  |  |  | -50 | mA |  | -50.0083 | mA | 0.0082979 | mA | 0.0057193 | mA
|  |  |  | 90 | mA |  | 90.0123 | mA | -0.010703 | mA | 0.008981 | mA
|  |  |  | -90 | mA |  | -90.0145 | mA | 0.014525 | mA | 0.0097691 | mA
|  |  |  | 100 | mA |  | 100.0138 | mA | -0.012105 | mA | 0.0099822 | mA
|  |  |  | -100 | mA |  | -100.0161 | mA | 0.01612 | mA | 0.0108 | mA
| 1 | A |  | 0.1 | A |  | 0.099984 | A | 1.7751e-05 | A | 1.0825e-05 | A
|  |  |  | -0.1 | A |  | -0.10001 | A | 1.4198e-05 | A | 1.0141e-05 | A
|  |  |  | 0.5 | A |  | 0.49999 | A | -1.6555e-06 | A | 0.00012743 | A
|  |  |  | -0.5 | A |  | -0.50001 | A | 4.3267e-05 | A | 0.00012677 | A
|  |  |  | 0.9 | A |  | 0.89997 | A | 3.1587e-06 | A | 0.00021382 | A
|  |  |  | -0.9 | A |  | -0.9 | A | 4.2176e-05 | A | 0.00021314 | A
|  |  |  | 1 | A |  | 0.99997 | A | -1.4528e-06 | A | 0.0002354 | A
|  |  |  | -1 | A |  | -0.99998 | A | 3.0943e-05 | A | 0.00023529 | A
Tegangan AC / AC Voltage
Rentang |  |  | Titik Ukur |  |  |  | Pembacaan Alat |  | Koreksi |  | Ketidakpastian
Range |  |  | Measurement Point |  |  |  | Instrument Reading |  | Correction |  | Uncertainty
| 10 | mV | 10 | mV* | 20 | Hz | 10 | mV | -0.00016155 | mV | 0.0057029 | mV
|  |  | 10 | mV* | 50 | Hz | 10 | mV | -0.00041446 | mV | 0.0056015 | mV
|  |  | 10 | mV* | 1 | kHz | 10 | mV | -0.00041446 | mV | 0.0056015 | mV
|  |  | 10 | mV* | 10 | kHz | 10 | mV | -0.0009925 | mV | 0.005674 | mV
|  |  | 10 | mV* | 20 | kHz | 10 | mV | -0.0052128 | mV | 0.0087477 | mV
| 100 | mV | 20 | mV* | 50 | kHz | 20 | mV | -0.012276 | mV | 0.018134 | mV
|  |  | 30 | mV | 100 | kHz | 30 | mV | -0.026265 | mV | 0.089601 | mV
|  |  | 50 | mV | 40 | Hz | 50 | mV | 0.00033009 | mV | 0.011684 | mV
|  |  | 100 | mV* | 80 | Hz | 100 | mV | 8.6035e-06 | mV | 0.014936 | mV
|  |  | 100 | mV | 20 | Hz | 100 | mV | 0.00059783 | mV | 0.024929 | mV
|  |  | 100 | mV* | 50 | Hz | 100 | mV | 8.6035e-06 | mV | 0.020635 | mV
|  |  | 100 | mV | 1 | kHz | 100 | mV | 8.6035e-06 | mV | 0.020635 | mV
|  |  | 100 | mV | 10 | kHz | 100 | mV | -0.012419 | mV | 0.046663 | mV
|  |  | 100 | mV | 20 | kHz | 100 | mV | 0.10445 | mV | 0.10716 | mV
|  |  | 100 | mV | 50 | kHz | 100 | mV | 0.10445 | mV | 0.11903 | mV
|  |  | 100 | mV | 100 | kHz | 100 | mV | -0.065749 | mV | 0.15354 | mV
| 1 | V | 0.1 | V | 20 | Hz | 0.1 | V | 5.9783e-07 | V | 2.4936e-05 | V
|  |  | 0.1 | V* | 50 | Hz | 0.1 | V | 8.6035e-09 | V | 2.0643e-05 | V
|  |  | 0.1 | V | 1 | kHz | 0.1 | V | 8.6035e-09 | V | 2.0643e-05 | V
|  |  | 0.1 | V | 10 | kHz | 0.1 | V | -1.2419e-05 | V | 4.6666e-05 | V
|  |  | 0.1 | V | 20 | kHz | 0.1 | V | 0.00010445 | V | 0.00010716 | V
|  |  | 0.1 | V | 50 | kHz | 0.1 | V | 0.00010445 | V | 0.00011903 | V
|  |  | 0.1 | V | 100 | kHz | 0.1 | V | -6.5749e-05 | V | 0.00015354 | V
|  |  | 0.3 | V* | 80 | Hz | 0.3 | V | 2.5641e-05 | V | 9.726e-05 | V
|  |  | 0.4 | V | 100 | Hz | 0.4 | V | 3.0924e-05 | V | 9.866e-05 | V
|  |  | 0.5 | V* | 160 | Hz | 0.5 | V | 3.6208e-05 | V | 0.00010045 | V
|  |  | 0.5 | V* | 200 | Hz | 0.5 | V | 3.6208e-05 | V | 0.00010045 | V
|  |  | 0.6 | V* | 400 | Hz | 0.6 | V | 4.1492e-05 | V | 0.0001023 | V
|  |  | 0.6 | V* | 500 | Hz | 0.6 | V | 4.1492e-05 | V | 0.0001023 | V
|  |  | 0.8 | V | 1 | kHz | 0.8 | V | 5.2059e-05 | V | 0.00010659 | V
|  |  | 0.8 | V* | 1.6 | kHz | 0.8 | V | 0.00010493 | V | 0.00012874 | V
|  |  | 1 | V | 20 | Hz | 1 | V | 7.1456e-05 | V | 0.00017563 | V
|  |  | 1 | V* | 50 | Hz | 1 | V | 6.2626e-05 | V | 0.00011161 | V
|  |  | 1 | V | 1 | kHz | 1 | V | 6.2626e-05 | V | 0.00011161 | V
|  |  | 1 | V | 10 | kHz | 1 | V | -5.8781e-05 | V | 0.00028502 | V
|  |  | 1 | V | 20 | kHz | 1 | V | -0.00051695 | V | 0.00081959 | V
|  |  | 1 | V | 50 | kHz | 1 | V | -0.00051695 | V | 0.00087936 | V
|  |  | 1 | V | 100 | kHz | 1 | V | -0.00046656 | V | 0.00097007 | V
| 10 | V | 1 | V | 20 | Hz | 1 | V | 7.1456e-05 | V | 0.00017572 | V
|  |  | 1 | V* | 50 | Hz | 1 | V | 6.2626e-05 | V | 0.00011175 | V
Tegangan AC / AC Voltage
Rentang |  |  | Titik Ukur |  |  |  | Pembacaan Alat |  | Koreksi |  | Ketidakpastian
Range |  |  | Measurement Point |  |  |  | Instrument Reading |  | Correction |  | Uncertainty
| 10 | V | 1 | V | 1 | kHz | 1 | V | 6.2626e-05 | V | 0.00011175 | V
|  |  | 1 | V* | 2 | kHz | 1 | V | 0.00013003 | V | 0.00013305 | V
|  |  | 1 | V* | 4 | kHz | 1 | V | 0.00013003 | V | 0.00013305 | V
|  |  | 1 | V* | 5 | kHz | 1 | V | 0.00013003 | V | 0.00013305 | V
|  |  | 1 | V | 10 | kHz | 1 | V | -5.8781e-05 | V | 0.00028508 | V
|  |  | 1 | V | 20 | kHz | 1 | V | -0.00051695 | V | 0.00081961 | V
|  |  | 1 | V | 50 | kHz | 1 | V | -0.00051695 | V | 0.00087937 | V
|  |  | 1 | V | 100 | kHz | 1 | V | -0.00046656 | V | 0.00097009 | V
|  |  | 5 | V | 20 | Hz | 5 | V | 0.00038142 | V | 0.0013852 | V
|  |  | 5 | V* | 50 | Hz | 5 | V | 0.00059717 | V | 0.0010015 | V
|  |  | 5 | V | 1 | kHz | 5 | V | 0.00059717 | V | 0.0010015 | V
|  |  | 5 | V | 10 | kHz | 5 | V | -0.00039531 | V | 0.0028068 | V
|  |  | 5 | V | 20 | kHz | 5 | V | -0.0009867 | V | 0.0081537 | V
|  |  | 5 | V | 50 | kHz | 5 | V | -0.0009867 | V | 0.0081702 | V
|  |  | 5 | V | 20 | kHz | 5 | V | -0.0009867 | V | 0.0081537 | V
|  |  | 10 | V | 20 | Hz | 10 | V | 0.00064346 | V | 0.0017386 | V
|  |  | 10 | V* | 50 | Hz | 10 | V | 0.00095745 | V | 0.0011 | V
|  |  | 10 | V | 1 | kHz | 10 | V | 0.00095745 | V | 0.0011 | V
|  |  | 10 | V | 10 | kHz | 10 | V | -0.00087322 | V | 0.0028434 | V
|  |  | 10 | V | 20 | kHz | 10 | V | -0.0021324 | V | 0.0081897 | V
|  |  | 10 | V | 50 | kHz | 10 | V | -0.0021324 | V | 0.0082294 | V
|  |  | 10 | V | 100 | kHz | 10 | V | -0.003822 | V | 0.0091757 | V
| 100 | V | 10 | V | 20 | Hz | 10 | V | 0.00064346 | V | 0.0017396 | V
|  |  | 10 | V* | 50 | Hz | 10 | V | 0.00095745 | V | 0.0011015 | V
|  |  | 10 | V | 1 | kHz | 10 | V | 0.00095745 | V | 0.0011015 | V
|  |  | 10 | V | 10 | kHz | 10 | V | -0.00087322 | V | 0.002844 | V
|  |  | 10 | V | 20 | kHz | 10 | V | -0.0021324 | V | 0.0081899 | V
|  |  | 10 | V | 50 | kHz | 10 | V | -0.0021324 | V | 0.0082296 | V
|  |  | 10 | V | 100 | kHz | 10 | V | -0.003822 | V | 0.0091758 | V
|  |  | 100 | V | 20 | Hz | 100 | V | -0.00028637 | V | 0.017394 | V
|  |  | 100 | V* | 50 | Hz | 100 | V | -0.0025353 | V | 0.011792 | V
|  |  | 100 | V | 1 | kHz | 100 | V | -0.0025353 | V | 0.011792 | V
|  |  | 100 | V | 10 | kHz | 100 | V | -0.029216 | V | 0.028761 | V
|  |  | 100 | V | 20 | kHz | 100 | V | -0.041907 | V | 0.082385 | V
|  |  | 100 | V | 50 | kHz | 100 | V | -0.041907 | V | 0.084474 | V
|  |  | 100 | V | 100 | kHz | 100 | V | -0.040392 | V | 0.14883 | V
| 1000 | V | 100 | V | 20 | Hz | 100 | V | -0.00028637 | V | 0.017403 | V
|  |  | 100 | V* | 50 | Hz | 100 | V | -0.0025353 | V | 0.011806 | V
|  |  | 100 | V | 1 | kHz | 100 | V | -0.0025353 | V | 0.011806 | V
|  |  | 500 | V* | 50 | Hz | 500 | V | 0.071289 | V | 0.075883 | V
Tegangan AC / AC Voltage
Rentang |  |  | Titik Ukur |  |  |  | Pembacaan Alat |  | Koreksi |  | Ketidakpastian
Range |  |  | Measurement Point |  |  |  | Instrument Reading |  | Correction |  | Uncertainty
| 1000 | V | 500 | V | 1 | kHz | 500 | V | 0.071289 | V | 0.075883 | V
|  |  | 900 | V* | 50 | Hz | 900 | V | 0.10968 | V | 0.12806 | V
|  |  | 900 | V | 1 | kHz | 900 | V | 0.10968 | V | 0.12806 | V
Arus AC / AC Current
Rentang |  |  | Titik Ukur |  |  |  | Pembacaan Alat |  | Koreksi |  | Ketidakpastian
Range |  |  | Measurement Point |  |  |  | Instrument Reading |  | Correction |  | Uncertainty
| 100 | µA | 10 | µA* | 20 | Hz | 10 | µA | -0.0024614 | µA | 0.12163 | µA
|  |  | 10 | µA* | 50 | Hz | 10 | µA | 0.00042726 | µA | 0.12075 | µA
|  |  | 10 | µA* | 1 | kHz | 10 | µA | 0.00042726 | µA | 0.12153 | µA
|  |  | 100 | µA* | 20 | Hz | 100 | µA | 0.0033687 | µA | 0.12827 | µA
|  |  | 100 | µA* | 50 | Hz | 100 | µA | 0.0033687 | µA | 0.12545 | µA
|  |  | 100 | µA* | 1 | kHz | 100 | µA | 0.0033687 | µA | 0.13207 | µA
| 1 | mA | 0.1 | mA* | 20 | Hz | 0.1 | mA | 2.2131e-06 | mA | 0.00012827 | mA
|  |  | 0.1 | mA* | 50 | Hz | 0.1 | mA | 3.3687e-06 | mA | 0.00012545 | mA
|  |  | 0.1 | mA* | 1 | kHz | 0.1 | mA | 3.3687e-06 | mA | 0.00013207 | mA
|  |  | 1 | mA | 20 | Hz | 1 | mA | 9.3378e-05 | mA | 0.00058261 | mA
|  |  | 1 | mA | 50 | Hz | 1 | mA | 8.6831e-05 | mA | 0.00055756 | mA
|  |  | 1 | mA | 1 | kHz | 1 | mA | 8.6831e-05 | mA | 0.00069139 | mA
|  |  | 1 | mA | 5 | kHz | 1 | mA | 6.5019e-05 | mA | 0.0023349 | mA
|  |  | 1 | mA | 10 | kHz | 1 | mA | -0.00016324 | mA | 0.0023404 | mA
| 10 | mA | 1 | mA | 20 | Hz | 1 | mA | 9.3378e-05 | mA | 0.00058264 | mA
|  |  | 1 | mA | 50 | Hz | 1 | mA | 8.6831e-05 | mA | 0.00055759 | mA
|  |  | 1 | mA | 1 | kHz | 1 | mA | 8.6831e-05 | mA | 0.00069141 | mA
|  |  | 1 | mA | 5 | kHz | 1 | mA | 6.5019e-05 | mA | 0.002335 | mA
|  |  | 1 | mA | 10 | kHz | 1 | mA | -0.00016324 | mA | 0.0023404 | mA
|  |  | 5 | mA | 20 | Hz | 5 | mA | 0.00071403 | mA | 0.0041696 | mA
|  |  | 5 | mA | 50 | Hz | 5 | mA | 0.00029322 | mA | 0.0040689 | mA
|  |  | 5 | mA | 1 | kHz | 5 | mA | 0.00029322 | mA | 0.0045134 | mA
|  |  | 5 | mA | 5 | kHz | 5 | mA | 0.00052722 | mA | 0.013085 | mA
|  |  | 5 | mA | 10 | kHz | 5 | mA | 0.00046212 | mA | 0.013182 | mA
|  |  | 10 | mA | 20 | Hz | 10 | mA | 0.0010566 | mA | 0.0058186 | mA
|  |  | 10 | mA | 50 | Hz | 10 | mA | 0.00073198 | mA | 0.0055757 | mA
|  |  | 10 | mA | 1 | kHz | 10 | mA | 0.00073198 | mA | 0.006255 | mA
|  |  | 10 | mA | 5 | kHz | 10 | mA | 0.00078076 | mA | 0.01955 | mA
|  |  | 10 | mA | 10 | kHz | 10 | mA | 0.00015365 | mA | 0.019615 | mA
| 100 | mA | 10 | mA | 20 | Hz | 10 | mA | 0.0010566 | mA | 0.0058189 | mA
|  |  | 10 | mA | 50 | Hz | 10 | mA | 0.00073198 | mA | 0.005576 | mA
|  |  | 10 | mA | 1 | kHz | 10 | mA | 0.00073198 | mA | 0.0062553 | mA
Arus AC / AC Current
Rentang |  |  | Titik Ukur |  |  |  | Pembacaan Alat |  | Koreksi |  | Ketidakpastian
Range |  |  | Measurement Point |  |  |  | Instrument Reading |  | Correction |  | Uncertainty
|  |  | 10 | mA | 5 | kHz | 10 | mA | 0.00078076 | mA | 0.01955 | mA
|  |  | 10 | mA | 10 | kHz | 10 | mA | 0.00015365 | mA | 0.019615 | mA
|  |  | 100 | mA | 20 | Hz | 100 | mA | 0.0070198 | mA | 0.058236 | mA
|  |  | 100 | mA | 50 | Hz | 100 | mA | 0.0074181 | mA | 0.0555 | mA
|  |  | 100 | mA | 1 | kHz | 100 | mA | 0.0074181 | mA | 0.060372 | mA
|  |  | 100 | mA | 5 | kHz | 100 | mA | 0.012396 | mA | 0.15042 | mA
|  |  | 100 | mA | 10 | kHz | 100 | mA | -0.002865 | mA | 0.15127 | mA
| 1 | A | 0.1 | A | 20 | Hz | 0.1 | A | 7.0198e-06 | A | 5.8239e-05 | A
|  |  | 0.1 | A | 50 | Hz | 0.1 | A | 7.4181e-06 | A | 5.5503e-05 | A
|  |  | 0.1 | A | 1 | kHz | 0.1 | A | 7.4181e-06 | A | 6.0375e-05 | A
|  |  | 0.1 | A | 5 | kHz | 0.1 | A | 1.2396e-05 | A | 0.00015042 | A
|  |  | 0.1 | A | 10 | kHz | 0.1 | A | -2.865e-06 | A | 0.00015127 | A
|  |  | 1 | A | 20 | Hz | 1 | A | -1.162e-05 | A | 0.00101 | A
|  |  | 1 | A | 50 | Hz | 1 | A | -2.5673e-06 | A | 0.0010101 | A
|  |  | 1 | A | 1 | kHz | 1 | A | -2.5673e-06 | A | 0.0011367 | A
|  |  | 1 | A | 5 | kHz | 1 | A | 0.00027853 | A | 0.0083378 | A
|  |  | 1 | A | 10 | kHz | 1 | A | 0.0011969 | A | 0.008344 | A
Resistansi / Resistance
Rentang |  |  | Titik Ukur |  |  |  | Pembacaan Alat |  | Koreksi |  | Ketidakpastian
Range |  |  | Measurement Point |  |  |  | Instrument Reading |  | Correction |  | Uncertainty
| 10 | Ω |  | 1 | Ω |  | 1 | Ω | 2.27e-05 | Ω | 1.4328e-05 | Ω
|  |  |  | 10 | Ω |  | 10 | Ω | 8e-06 | Ω | 0.00011802 | Ω
| 100 | Ω |  | 10 | Ω |  | 10 | Ω | 8e-06 | Ω | 0.00011802 | Ω
|  |  |  | 100 | Ω |  | 100 | Ω | 0.0005 | Ω | 0.0010081 | Ω
| 1 | kΩ |  | 0.1 | kΩ |  | 0.1 | kΩ | 5e-07 | kΩ | 1.0097e-06 | kΩ
|  |  |  | 1 | kΩ |  | 1 | kΩ | 5.71e-05 | kΩ | 5.8378e-05 | kΩ
| 10 | kΩ |  | 1 | kΩ |  | 1 | kΩ | 5.71e-05 | kΩ | 5.8381e-05 | kΩ
|  |  |  | 10 | kΩ |  | 10 | kΩ | 5.4e-05 | kΩ | 6.2153e-05 | kΩ
| 100 | kΩ |  | 10 | kΩ |  | 10 | kΩ | 5.4e-05 | kΩ | 6.2682e-05 | kΩ
|  |  |  | 100 | kΩ |  | 100 | kΩ | 0.00085 | kΩ | 0.0011089 | kΩ
| 1 | MΩ |  | 0.1 | MΩ |  | 0.1 | MΩ | 8.5e-07 | MΩ | 1.1104e-06 | MΩ
|  |  |  | 1 | MΩ |  | 1 | MΩ | 4e-06 | MΩ | 1.51e-05 | MΩ
| 10 | MΩ |  | 1 | MΩ |  | 1 | MΩ | 4e-06 | MΩ | 1.5111e-05 | MΩ
|  |  |  | 10 | MΩ |  | 10 | MΩ | 8e-05 | MΩ | 0.00026608 | MΩ
| 100 | MΩ |  | 10 | MΩ |  | 10 | MΩ | 8e-05 | MΩ | 0.00026615 | MΩ
|  |  |  | 100 | MΩ |  | 100 | MΩ | 0.0131 | MΩ | 0.0038678 | MΩ
| 1 | GΩ |  | 0.1 | GΩ |  | 0.1 | GΩ | 1.31e-05 | GΩ | 3.8682e-06 | GΩ
|  |  |  | 1 | GΩ |  | 1 | GΩ | 0.001067 | GΩ | 5.7469e-05 | GΩ
Catatan / Notes :
- | Hasil kalibrasi ini diperoleh berdasarkan prosedur kalibrasi I.ME.1.03 untuk tegangan DC, I.ME.3.04
untuk arus DC, I.ME.5.05 untuk tegangan AC, I.ME.6.03 untuk arus AC, dan I.ME.2.05 untuk
resistansi dengan menggunakan alat standar yang tertelusur ke SI melalui SNSU-BSN.
This calibration result was acquired based on the procedure of I.ME.1.03 for DC voltage , I.ME.3.04 for
DC current, I.ME.5.05 for AC voltage, I.ME.6.03 for AC current, and I.ME.2.05 for resistance using the standard
instrument that is traceable to SI through SNSU-BSN.
- | Alat standar yang digunakan adalah Multifunction Calibrator F.5730A (SN.4978506), Standard
Resistor F.742A-1 (SN.9411004), Standard Resistor F.742A-10 (SN.9411003), Standard Resistor
F.742A-100 (SN.9420003), Standard Resistor F.742A-1k (SN.9405005), Standard Resistor
F.742A-10k (SN.9420001), Standard Resistor F.742A-100k (SN.9411002), Standard Resistor
F.742A-1M (SN.9411001), Standard Resistor F.742A-10M (SN.9411011), Standard Resistor
G.9336-100M (SN.69073), dan Standard Resistor G.9336-1G (SN.68376).
The standard instrument used were Multifunction Calibrator F.5730A (SN.4978506), Standard Resistor F.742A-1
(SN.9411004), Standard Resistor F.742A-10 (SN.9411003), Standard Resistor F.742A-100 (SN.9420003),
Standard Resistor F.742A-1k (SN.9405005), Standard Resistor F.742A-10k (SN.9420001), Standard
Resistor F.742A-100k (SN.9411002), Standard Resistor F.742A-1M (SN.9411001), Standard Resistor
F.742A-10M (SN.9411011), Standard Resistor G.9336-100M (SN.69073), Standard Resistor
G.9336-1G (SN.68376)
- | Pengukuran resistansi pada rentang 1 Ω sampai dengan 100 kΩ menggunakan metode 4-kawat,
sedangkan pada rentang yang lain menggunakan metode 2-kawat.
The measurement of resistances at the range of 1 Ω to 100 kΩ used 4-wire method, while the others
used 2-wire method.
- | Ketidakpastian pengukuran dihitung dengan tingkat kepercayaan tidak kurang dari 95% dan faktor
cakupan k = 2.
The uncertainty of measurement was calculated with the confidence level not less than 95 % and
coverage factor of k = 2.
- | Hasil kalibrasi yang ditandai bintang (*) tidak tercakup dalam ruang lingkup akreditasi KAN.
Calibration results marked by asterisk (*) are not covered by KAN accreditation.
Dikalibrasi oleh/Calibrated by | [6 empty] | : Hayati Amalia, S.T.
Diperiksa oleh/Checked by | [6 empty] | : Agah Faisal, M.Sc & Lukluk Khairiyati, M.T.
(Penyelia/Supervisor)
: Agah Faisal, M.Sc.
(Ka. Lab SNSU Kelistrikan/Head of NMS for Electricity Laboratory)
= = = = Akhir dari Sertifikat / End of Certificate = = = =